|  |  |  | Test Report issued under the responsibility of: |
| --- | --- | --- | --- |
| ### TEST REPORT IEC 61010-1 Safety requirements for electrical equipment for measurement, control, and laboratory use Part 1: General requirements |  |  |  |  |
| Report Number. : |  |  |  |  |
| Date of issue : |  |  |  |  |
| Total number of pages : |  |  |  |  |
| Name of Testing Laboratory preparing the Report : |  |  |  |  |
| Applicant’s name : | Gener8 LLC |  |  |  |
| Address : | 5975 Castle Creek Pkwy N Dr, Suite 250, Indianapolis, IN 46250-2781, UNITED STATES |  |  |  |
| Test specification: |  |  |  |  |
| Standard : | IEC 61010-1:2010, IEC 61010-1:2010/AMD1:2016 |  |  |  |
| Test procedure : | CB Scheme |  |  |  |
| Non-standard test method : | N/A |  |  |  |
| TRF template used : | IECEE OD-2020-F1:2020, Ed.1.3 |  |  |  |
| Test Report Form No. : | IEC61010_1P |  |  |  |
| Test Report Form(s) Originator : | VDE Prüf- und Zertifizierungsinstitut GmbH |  |  |  |
| Master TRF : | 2021-04-12 |  |  |  |
| Copyright © 2021 IEC System of Conformity Assessment Schemes for Electrotechnical Equipment and Components (IECEE System). All rights reserved. This publication may be reproduced in whole or in part for non-commercial purposes as long as the IECEE is acknowledged as copyright owner and source of the material. IECEE takes no responsibility for and will not assume liability for damages resulting from the reader's interpretation of the reproduced material due to its placement and context. If this Test Report Form is used by non-IECEE members, the IECEE/IEC logo and the reference to the CB Scheme procedure shall be removed. This report is not valid as a CB Test Report unless signed by an approved IECEE Testing Laboratory and appended to a CB Test Certificate issued by an NCB in accordance with IECEE 02. |  |  |  |  |
| General disclaimer: |  |  |  |  |
| The test results presented in this report relate only to the object tested. This report shall not be reproduced, except in full, without the written approval of the Issuing NCB. The authenticity of this Test Report and its contents can be verified by contacting the NCB, responsible for this Test Report. |  |  |  |  |

| Test item description : |  | CellFE Low Throughput Instrument |  |  |
| --- | --- | --- | --- | --- |
| Trade Mark : |  | Gener8 LLC |  |  |
| Manufacturer : |  | Gener8 LLC |  |  |
| Model/Type reference : |  | CellIFE Low Throughput Instrument |  |  |
| Ratings : |  | 100-240VAC, 5.9-2.7A, 50/60Hz; 24VDC output |  |  |
| Responsible Testing Laboratory (as applicable), testing procedure and testing location(s): |  |  |  |  |
|  | CB Testing Laboratory: |  |  |  |
| Testing location/ address : |  |  |  |  |
| Tested by (name, function, signature) : |  |  |  |  |
| Approved by (name, function, signature) : |  |  |  |  |
|  | Testing procedure: CTF Stage 1: |  |  |  |
| Testing location/ address : |  |  |  |  |
| Tested by (name, function, signature) : |  |  |  |  |
| Approved by (name, function, signature) : |  |  |  |  |
|  | Testing procedure: CTF Stage 2: |  |  |  |
| Testing location/ address : |  |  |  |  |
| Tested by (name + signature) : |  |  |  |  |
| Witnessed by (name, function, signature) : |  |  |  |  |
| Approved by (name, function, signature) : |  |  |  |  |
|  | Testing procedure: CTF Stage 3: |  |  |  |
|  | Testing procedure: CTF Stage 4: |  |  |  |
| Testing location/ address : |  |  |  |  |
| Tested by (name, function, signature) : |  |  |  |  |
| Witnessed by (name, function, signature) : |  |  |  |  |
| Approved by (name, function, signature) : |  |  |  |  |
| Supervised by (name, function, signature) : |  |  |  |  |

| List of Attachments (including a total number of pages in each attachment) |  |  |
| --- | --- | --- |
| Document No. | Documents included / attached to this report (description) | Page No. |

| Documents referenced by this report (available on request): |  |  |
| --- | --- | --- |
| Document Name or No. | Documents description | Page No. |

| Summary of testing: |  |
| --- | --- |
| Clause | Comment |

| Test Report History: This report may consist of more than one report and is only valid with additional or previous issued reports: |  |
| --- | --- |
| Report Ref. No. | Item |
| Tests performed (name of test and test clause): | Testing location: |
| Summary of compliance with National Differences (List of countries addressed): The product fulfils the requirements of: US |  |
| Statement concerning the uncertainty of the measurement systems used for the tests (may be required by the product standard or client) Internal procedure used for type testing through which traceability of the measuring uncertainty has been established: Procedure number, issue date and title: Calculations leading to the reported values are on file with the NCB and testing laboratory that conducted the testing. Statement not required by the standard used for type testing (Note: When IEC or ISO standard requires a statement concerning the uncertainty of the measurement systems used for tests, this should be reported above. The informative text in parenthesis should be delete in both cases after selecting the applicable option) |  |

Copy of marking plate: The artwork below may be only a draft. The use of certification marks on a product must be authorized by the respective NCBs that own these marks.

| Test item particulars: |  |  |
| --- | --- | --- |
| Type of item : |  | Laboratory |
| Description of equipment function : |  | Automated pneumatic system for lab air flow and cartridge handling. |
| Connection to mains supply : |  | Permanent |
| Overvoltage category : |  | II |
| Pollution degree : |  | Pollution degree 2 |
| Means of protection : |  | Class I (PE connected) |
| Environmental conditions : |  | Normal |
| For use in wet locations : |  | No |
| Equipment mobility : |  | Floor standing |
| Operating conditions : |  | Continuous |
| Overall size of equipment (W x D x H) : |  | Overall size not specified in provided documents/images. |
| Mass of equipment (kg) : |  | Not available in provided documents or images. |
| Marked degree of protection to IEC 60529 : |  | No IPxx rating marking found. |
| Possible test case verdicts: |  |  |
| - Test case does not apply to the test object : |  | N/A (Not Applicable) |
| - Test object does meet the requirement : |  | P (Pass) |
| - Test object does not meet the requirement : |  | F (Fail) |
| Testing: |  |  |
| Date of receipt of test item : |  |  |
| Date (s) of performance of tests : |  |  |
| General remarks: |  |  |
| The test results presented in this report relate only to the object tested. This report shall not be reproduced, except in full, without the written approval of the issuing testing laboratory. "(see Enclosure #)" refers to additional information appended to the report. "(see Form A.xx)" refers to a Table appended to the report. Bottom lines for measurement Tables Forms A.xx are optional if used as record. Throughout this report a comma / point is used as the decimal separator. |  |  |
| Manufacturer’s Declaration per sub-clause 4.2.5 of IECEE 02: |  |  |
| The application for obtaining a CB Test Certificate includes more than one factory location and a declaration from the manufacturer stating that the sample(s) submitted for evaluation is (are) representative of the products from each factory has been provided : |  | Yes Not applicable |
| When differences exist; they shall be identified in the general product information section. |  |  |
| Name and address of factory (ies) : | Gener8 LLC, 2560 Junction Ave, San Jose, CA 95134, UNITED STATES |  |

General product information and other remarks: Description of unit: The equipment is a CellFE Low Throughput Instrument, integrating multiple electronic and pneumatic subsystems. It features dual cameras, IR LED rings, sensors, valves, and a Jetson TX2 for processing. The system operates with a 110/220VAC input, converting to 15VDC and 24VDC for internal circuits, and includes safety features like fuses and pressure relief valves. It interfaces with peripherals such as a touchscreen, barcode scanner, and NFC reader, and is powered by both AC mains and a LiPo battery. The enclosure is metal and grounded for safety.

Description of model differences: N/A- one model

Description of special features: This product includes pneumatic systems with high-pressure air circuits, multiple pressure sensors, electronic regulators, and pressure relief safety valves rated up to 125 PSI. It also integrates various flow sensors for both low and high flow rates, and uses a 3S LiPo battery for backup power. The system operates with multiple DC voltages and includes a metal enclosure with chassis grounding. There are no high voltage (>1000 V) circuits, but the presence of high-pressure pneumatic components and multiple regulated power supplies are notable special features.

| IEC 61010-1 |  |  |  |
| --- | --- | --- | --- |
| Clause | Requirement + Test | Result - Remark | Verdict |
| 4 | TESTS |  |  |
| 4.4 | Testing in single fault conditions |  |  |
| 4.4.1 | Fault tests | (see Form A.1) |  |
| 4.4.2 | Application of single fault conditions |  |  |
| 4.4.2.1 | single fault conditions not covered by 4.4.2.2 to 4.4.2.14 | (see Form A.1) |  |
| 4.4.2.2 | protective impedance | TBD-Info available |  |
| 4.4.2.3 | protective conductor | (see Form A.6) |  |
| 4.4.2.4 | Equipment or parts for short-term or intermittent operation | TBD-Info available |  |
| 4.4.2.5 | Motors | TBD-Info available | — |
|  | – stopped while fully energized | TBD-Info available |  |
|  | – prevented from starting | TBD-Info not available |  |
|  | – one phase interrupted (multi-phase) | TBD-Info not available |  |
| 4.4.2.6 | Capacitors | TBD-Info available |  |
| 4.4.2.7 | mains transformers | TBD-Info available |  |
| 4.4.2.7.2 | Short circuit | (see Form A.39) |  |
| 4.4.2.7.3 | Overload | (see Forms A.26B and A.40) |  |
| 4.4.2.8 | Outputs | TBD-Info available |  |
| 4.4.2.9 | Equipment for more than one supply | TBD-Info available |  |
| 4.4.2.10 | Cooling | (see Form A.26A) | — |
|  | – air holes closed | TBD-Info not available |  |
|  | – fans stopped | TBD-Info not available |  |
|  | – coolant stopped | TBD-Info not available |  |
|  | – loss of cooling liquid | TBD-Info not available |  |
| 4.4.2.11 | Heating devices | TBD-Info available | — |
|  | – timer overridden | TBD-Info available |  |
|  | – temperature controller overridden | TBD-Info available |  |
| 4.4.2.12 | Insulation between circuits and parts | TBD-Info available |  |
| 4.4.2.13 | Interlocks | TBD-Info not available |  |
| 4.4.2.14 | Voltage selectors | TBD-Info available |  |
| 4.4.3 | Duration of tests | (see Form A.1) | — |
| 4.4.4 | Conformity after application of fault conditions | (see Forms A.1, A.6 and A.18) |  |
| 5 | MARKING AND DOCUMENTATION |  |  |
| 5.1 | Marking |  |  |
| 5.1.1 | General |  |  |
|  | Required equipment markings | Critical marking label image not provided for verification. | — |
|  | – Visible from the exterior; or | Required exterior markings not visible in provided images. | N/A |
|  | – Visible after removing cover or opening door | Critical marking label images not provided for verification. | N/A |
|  | – Visible after removal from a rack or panel | Critical marking visibility evidence not available in images. | N/A |
|  | Not put on parts which can be removed by an operator | No evidence of operator-removable parts with markings. | P |
|  | Letter symbols (IEC 60027) used |  |  |
|  | Graphic symbols of Table 1 used | Critical marking symbol information not available in documentation. | N/A |
| 5.1.2 | Identification |  | P |
|  | Equipment is identified by: |  | — |
|  | - Manufacturer’s or supplier’s name or trademark | Gener8 name visible on purchase order document. | P |
|  | - Model number, name or other means | Model name 'CellIFE Low Throughput Instrument' identified. | P |
|  | Manufacturing location identified | Manufacturing location identified as United States. | P |
| 5.1.3 | mains supply |  |  |
|  | Equipment is marked as follows: |  | — |
|  | - Nature of supply: | Mains supply is 100-240VAC, 50/60Hz. | — |
|  | - a.c. rated mains frequency or range of frequencies : | Rated mains frequency is 50/60Hz. | — |
|  | - d.c. with symbol 1 : | No evidence of d.c. symbol 1 for mains supply. | — |
|  | - rated supply voltage(s) or range : | Rated supply voltage: 100-240VAC, 50/60Hz. | — |
|  | - Max. rated power (W or VA) or input current : |  | — |
|  | The marked value not less than 90 % of the maximum value | (see Form A.2) |  |
|  | If more than one voltage range: | TBD-Info available | — |
|  | Separate values marked; or | TBD-Info available |  |
|  | Values differ by less than 20 % | (see Form A.2) |  |
|  | - operator-set for different rated supply voltages: | Operator-set for supply voltages not evident in documentation. | — |
|  | Indicates the equipment set voltage | Operator-set voltage indication not evidenced in documentation. | N/A |
|  | portable equipment indication is visible from the exterior | Critical marking label image not provided for verification. | N/A |
|  | Changing the setting changes the indication | Critical information on operator-set voltage indication missing. | N/A |
|  | - Accessory mains socket-outlets accepting standard mains plugs are marked: | No accessory mains socket-outlets observed or marked. | — |
|  | With the voltage if it is different from the mains supply voltage : | No accessory mains socket-outlets identified in documentation. | — |
|  | For use only with specific equipment | No accessory mains socket-outlets identified in documentation. | N/A |
|  | If not marked for specific equipment it is marked with: | No accessory mains socket-outlets identified in documentation. | — |
|  | The maximum rated current or power; or | No accessory mains socket-outlets identified in documentation. | N/A |
|  | Symbol 14 with full details in the documentation | No accessory mains socket-outlets identified in documentation. | N/A |
| 5.1.4 | Fuses | Fuse present on AC input per block diagram. |  |
|  | operator replaceable fuse marking (see also 5.4.5) : | Critical marking information for fuses not available in files. | — |
| 5.1.5 | Terminals, connections and operating devices |  |  |
| 5.1.5.1 | General |  |  |
|  | Where necessary for safety, indication of purpose of terminals, connectors, controls and indicators marked | Purpose of terminals and connectors is clearly marked. | P |
|  | If insufficient space, symbol 14 used | Critical information on symbol 14 usage not available. | N/A |
|  | Push-buttons and actuators of emergency stop devices and indicators: |  | — |
|  | – used only to indicate a warning of danger; or | No emergency stop devices or warning indicators identified. | N/A |
|  | – the need for urgent action | Critical information on emergency stop devices not available. | N/A |
|  | – coloured red | Critical information on emergency stop device color not available. | N/A |
|  | – coded as specified in IEC 60073 |  |  |
|  | Supplementary means of coding provided, if meaning of colour relates (see IEC 60073): |  | — |
|  | – to safety of persons; or | Critical information on emergency stop devices is not available. | N/A |
|  | – safety of the environment | Critical information on emergency stop devices not available. | N/A |
| 5.1.5.2 | terminals |  | — |
|  | mains supply terminal identified | Mains supply terminal clearly identified as L, N, GND. | P |
|  | Other terminal marking: |  | — |
|  | - functional earth terminals marked with symbol 5 | Critical marking label image for symbol 5 not available. | N/A |
|  | - protective conductor terminals: |  | — |
|  | Symbol 6 is placed close to or on the terminal; or | Critical marking symbol information not available in documentation. | P |
|  | Part of appliance inlet | Critical marking information for protective conductor terminal missing. | P |
|  | - terminals of circuits (symbol 7 used) | Chassis ground marked with symbol 7 in block diagram. | P |
|  | - hazardous live terminals supplied from the interior | Marking of hazardous live terminals not evidenced. | N/A |
|  | Standard mains socket outlet used; or | Marking and socket outlet evidence not available in files. | P |
|  | ratings marked; or | Hazardous live terminals ratings not explicitly marked in images. | N/A |
|  | Symbol 14 used | Critical marking symbol evidence not available in documentation. | N/A |
| 5.1.6 | Switches and circuit-breakers |  |  |
|  | If disconnecting device, off position clearly marked | Critical marking information for OFF position not available. | N/A |
|  | If push-button used as power supply switch: |  | — |
|  | – Symbol 9 and 15 used for on-position | Critical marking label image not provided for verification. | N/A |
|  | – Symbol 10 and 16 used for off-position | Critical information on push-button symbols is not available. | N/A |
|  | – Pair of symbols 9, 15 and 10, 16 close together | Critical information on push-button symbols not available. | N/A |
| 5.1.7 | Equipment protected by double insulation or reinforced insulation |  | N/A |
|  | Protected throughout (symbol 11 used) | Symbol 11 marking evidence not found in documentation. | N/A |
|  | Only partially protected (symbol 11 not used) | Critical marking label image not provided for verification. | N/A |
| 5.1.8 | Field-wiring terminal boxes |  |  |
|  | If terminal or enclosure exceeds 60 C: | (see Form A.26A) | — |
|  | Cable temperature rating marked : | Required cable temperature rating marking not evidenced. | — |
|  | Marking visible before and during connection or beside terminal | Critical marking visibility information not available in documentation. | N/A |
| 5.2 | Warning markings |  |  |
|  | Visible when ready for normal use | Critical marking label images not provided for verification. | N/A |
|  | Are near or on applicable parts | Warning marking placement evidence not available in images. | N/A |
|  | Symbols and text correct dimensions and colour: |  | — |
|  | - Symbols min 2,75 mm and text 1,5 mm high and contrasting in colour with background | Critical marking size and contrast information not available. | N/A |
|  | - Symbols and text moulded, stamped or engraved in material min. 2,0 mm high and | Critical marking dimension and color evidence not available. | N/A |
|  | 0,5 mm depth or raised if not contrasting in colour | Critical marking label details not available in documentation. | N/A |
|  | If necessary marked with symbol 14, or | Critical marking label images not provided for verification. | N/A |
|  | Additional symbols such as symbol 12, 13 or 17 used to indicate the nature of hazard | Critical marking symbol evidence not available in documentation. | N/A |
|  | Statement to place equipment in a safe state before access by using a tool to hazardous parts is permitted | Critical marking statement evidence not available in documentation. | N/A |
| 5.3 | Durability of markings |  |  |
|  | The required markings remain clear and legible in normal use | (see Form A.3) |  |
| 5.4 | Documentation |  |  |
| 5.4.1 | General | Required documentation is not fully available for review. |  |
|  | Equipment is accompanied by documentation for safety purposes for operator or responsible body | Critical information on safety documentation not available. | N/A |
|  | Safety documentation for service personnel authorized by the manufacturer | Critical information on service documentation not available. | N/A |
|  | Documentation necessary for safe operation is provided in printed media or | Critical documentation for safe operation not evidenced. | N/A |
|  | in electronic media if available at any time | Documentation in electronic form is available anytime. | P |
|  | Documentation includes: |  | — |
|  | - Intended use | Intended use not stated in available documentation. | N/A |
|  | - Technical specification | Technical specification details are not included in documentation. | P |
|  | - Name and address of manufacturer or supplier | Manufacturer: Gener8 LLC, 2560 Junction Ave, San Jose, CA 95134 | P |
|  | - Information specified in 5.4.2 to 5.4.6 | Insulation types and voltages are clearly documented. | P |
|  | - Information to mitigate residual risk (see also subclause 17) | Residual risk mitigation info not found in documentation. | N/A |
|  | - Accessories for safe operation of the equipment specified | No accessories for safe operation specified in documentation. | N/A |
|  | - Guidance provided to check correct function of the equipment, if incorrect reading may cause a hazard from harmful or corrosive substances of hazardous live parts | Critical information not available in documentation. | N/A |
|  | - Instructions for lifting and carrying | Lifting and carrying instructions not found in documentation. | N/A |
|  | Warning statements and a clear explanation of warning symbols: | Warning statements and symbol explanations not found. | — |
|  | – provided in the documentation; or | Warning statements and symbol explanations not found. | N/A |
|  | – information is marked on the equipment | Warning statements and symbols not evidenced in files. | N/A |
| 5.4.2 | Equipment ratings |  |  |
|  | Documentation includes: |  | — |
|  | - Supply voltage or voltage range : | Supply voltage is 100-240VAC, 24VDC output. | — |
|  | Frequency or frequency range : | Frequency is 50/60Hz per block diagram evidence. | — |
|  | Power or current rating : | Power rating: 100-240VAC, 5.9-2.7A, 50/60Hz, 24VDC output | — |
|  | - Description of all input and output connections in accordance to 6.6.1 a) |  |  |
|  | - rating of insulation of external circuits in accordance to 6.6.1 b) |  |  |
|  | - Statement of the range of environmental conditions (refer to 1.4): | Environmental conditions statement not found in documentation. | — |
|  | 1) indoor or outdoor use, | Indoor use only; no outdoor rating evidence found. | P |
|  | 2) altitude, | Altitude range information not available in documentation. | N/A |
|  | 3) temperature, | Temperature range information not available in documentation. | N/A |
|  | 4) relative humidity, | Relative humidity range not specified in documentation. | N/A |
|  | 5) mains supply voltage fluctuations, | Mains supply fluctuation range: 100-240VAC specified. | P |
|  | 6) overvoltage category, | Overvoltage category II is specified in documentation. | P |
|  | 7) wet location, if applicable, | No evidence of wet location use found. | N/A |
|  | 8) pollution degree of the intended environment | Pollution degree 2 is specified in TRF documentation. | P |
|  | - Degree of ingress protection (IEC 60529) |  |  |
|  | - If impact rating less than 5 J: |  | — |
|  | IK code in accordance to IEC 62262 marked; or | No IK code marking evidence found in documentation. | N/A |
|  | symbol 14 of Table 1 marked, with | Critical marking symbol evidence not available in images. | N/A |
|  | rated energy level and test method stated | Impact energy rating and test method not documented. | N/A |
| 5.4.3 | Equipment installation |  |  |
|  | Documentation includes instructions for: |  | — |
|  | - Assembly, location and mounting requirements | Critical installation instructions not found in documentation. | N/A |
|  | - Instructions for protective earthing | Installation instructions for protective earthing not found. | N/A |
|  | - Connections to supply | Instructions for supply connections not found in documentation. | P |
|  | - permanently connected equipment: | Installation instructions for permanent connection not found. | — |
|  | 1) Supply wiring requirements | Installation instructions for supply wiring not found. | N/A |
|  | 2) If external switch or circuit-breaker, requirements and location recommendation | Critical information on external switches not available. | N/A |
|  | - Ventilation requirements | Ventilation instructions not found in documentation provided. | N/A |
|  | - Safety characteristics for special external services (e. g. maximum and minimum temperature, pressure, flow of air, cooling liquid) | Critical installation safety characteristics not documented. | N/A |
|  | - Instructions relating to sound level | Sound level installation instructions not found in documentation. | N/A |
| 5.4.4 | Equipment operation |  |  |
|  | Instructions for use include: |  | — |
|  | - Identification and description of operating controls | Critical information not available in the documentation. | N/A |
|  | - Positioning for disconnection | Critical information not available in the documentation. | N/A |
|  | - Instructions for interconnection to accessories or other equipment | Critical information not available in the documentation. | N/A |
|  | - Specification of intermittent operation limits | Intermittent operation limits not specified in documentation. | N/A |
|  | - Explanation of symbols used | Critical information on symbol explanations is not available. | N/A |
|  | - Replacement of consumable materials | Critical information not available in the documentation. | N/A |
|  | - Cleaning and decontamination | Cleaning and decontamination instructions not found. | N/A |
|  | - Listing of any poisonous or injurious gases and quantities | No information on poisonous or injurious gases found. | N/A |
|  | - risk reduction procedures relating to flammable liquids (see 9.5 c) | No flammable liquid risk-reduction procedures found. | N/A |
|  | - risk reduction procedures relating burn from surfaces permitted to exceed limits of 10.1 | Critical information not available in documentation. | N/A |
|  | Additional precautions for IEC 60950 conforming equipment in regard to moistures and liquids |  |  |
|  | A statement about protection impairment if used in a manner not specified by the manufacturer | Critical information not available in the documentation. | N/A |
| 5.4.5 | Equipment maintenance and service |  |  |
|  | Instructions for responsible body include: | Maintenance and service instructions not found in documentation. | — |
|  | Instructions sufficient in detail permitting safe maintenance and inspection and continued safety: | Critical maintenance instructions not found in documentation. | — |
|  | Instruction against the use of detachable mains supply cord with inadequate rating | Critical information not available in the documentation. | N/A |
|  | Specific battery type of user replaceable batteries | Exact battery type not specified in instructions. | N/A |
|  | Any manufacturer specified parts | Manufacturer-specified parts for service not documented. | N/A |
|  | rating and characteristics of fuses | Fuse rating and characteristics not documented. | N/A |
|  | Instructions include following subjects permitting safe servicing and continued safety: | Maintenance and service instructions not found in documentation. | — |
|  | - Product specific risks may affect service personnel | Critical maintenance instructions not found in documentation. | N/A |
|  | - Protective measures for these risks | Critical information on service instructions not available. | N/A |
|  | - Verification of the safe state after repair | Critical maintenance instructions not found in documentation. | N/A |
| 5.4.6 | Integration into systems or effects resulting from special conditions |  | P |
|  | Aspects described in documentation | Critical information on maintenance and servicing is not available. | P |
| 6 | PROTECTION AGAINST ELECTRIC SHOCK |  |  |
| 6.1 | General | (see Forms A.14 and A.15) |  |
| 6.1.1 | Requirements | TBD-Info available |  |
|  | Protection against electric shock maintained in normal condition and single fault condition | TBD-Info available |  |
|  | accessible parts not hazardous live | TBD-Info available |  |
|  | Voltage, current, charge or energy below the limits in normal condition and in single fault condition between: | TBD-Info available | — |
|  | accessible parts and earth | TBD-Info available |  |
|  | two accessible parts on same piece of the equipment within a distance of 1,8 m | TBD-Info available |  |
|  | Conformity is checked by the determination of 6.2 and 6.3 followed by the tests of 6.4 to 6.11 |  |  |
| 6.1.2 | Exceptions | TBD-Info available |  |
|  | Following hazardous live parts may be accessible to an operator: | TBD-Info available | — |
|  | - parts of lamps and lamp sockets after lamp removal | TBD-Info available |  |
|  | - parts to be replaced by operator only by the use of tool and warning marking | TBD-Info available |  |
|  | Those parts not hazardous live 10 s after interruption of supply | (see Form A.5 ) |  |
|  | Capacitance test if charge is received from internal capacitor | (see Forms A.4 and A.5) |  |
| 6.2 | Determination of accessible parts | (see Form A.4) |  |
| 6.2.1 | General |  |  |
|  | Unless obviously determination of accessible parts as specified in 6.2.2 to 6.2.4 |  |  |
| 6.2.2 | Examination |  |  |
|  | – with jointed test finger (as specified B.2) | TBD-Info available |  |
|  | – with rigid test finger (as specified B.1) and a force of 10 N | TBD-Info not available |  |
| 6.2.3 | Openings above parts that are hazardous live |  |  |
|  | – test pin with length of 100 mm and 4 mm in diameter applied | TBD-Info available |  |
| 6.2.4 | Openings for pre-set controls |  |  |
|  | – test pin with length of 100 mm and 3 mm in diameter applied | TBD-Info available |  |
| 6.3 | Limit values for accessible parts |  |  |
| 6.3.1 | Levels in normal condition | (see Form A.5) |  |
|  | - Voltage limits less than 30 V r.m.s. and 42,4 V peak or 60 V d.c. | TBD-Info available |  |
|  | for wet locations voltage limits less than 16 V r.m.s. and 22,6 V peak or 35 V d.c. | TBD-Info available |  |
|  | Voltages are not hazardous live the levels of: |  | — |
|  | - Current less than 0,5 mA r.m.s. for sinusoidal, 0,7 mA peak non-sinusoidal or mixed frequencies or 2 mA d.c. when measured with measuring circuit A.1 or A.2 if less than 100 Hz | TBD-Info available |  |
|  | for wet locations measuring circuit A.4 used | TBD-Info available |  |
|  | 70 mA r.m.s. when measured with circuit A.3 for higher frequencies | TBD-Info available |  |
|  | - Levels of capacitive charge or energy less: |  | — |
|  | 1) 45 µC for voltages up to 15 kV peak or d.c. or line A of Figure 3 | TBD-Info available |  |
|  | 2) 350 mJ stored energy for voltages above 15 kV peak or d.c. | TBD-Info available |  |
| 6.3.2 | Levels in single fault condition | (see Form A.6) |  |
|  | - Voltage limits less than 50 V r.m.s. and 70 V peak or 120 V d.c. | TBD-Info available |  |
|  | for wet locations voltage limits less than 33 V r.m.s. and 46,7 V peak or 70 V d.c. | TBD-Info available |  |
|  | Voltages are not hazardous live the levels of: |  | — |
|  | - Current less than 3,5 mA r.m.s. for sinusoidal, 5 mA peak non-sinusoidal or mixed frequencies or 15 mA d.c. when measured with measuring circuit A.1 or A.2 if less than 100 Hz | TBD-Info available |  |
|  | for wet locations measuring circuit A.4 used | TBD-Info available |  |
|  | 500 mA r.m.s. when measured with circuit A.3 for higher frequencies | TBD-Info available |  |
|  | - Levels of capacitive charge or energy less line B of Figure 3 | TBD-Info available |  |
| 6.4 | Primary means of protection |  |  |
| 6.4.1 | General |  |  |
|  | accessible parts prevented from being hazardous live by one or more of following means: |  | — |
|  | - enclosures or protective barriers (see 6.4.2) |  |  |
|  | - basic insulation (see 6.4.3) |  |  |
|  | - Impedance (see 6.4.4) |  |  |
| 6.4.2 | enclosures or protective barriers | (see Forms A.15 and A.16) |  |
|  | – meet rigidity requirements of 8.1 |  |  |
|  | – meet requirements for basic insulation, if protection is provided by insulation | TBD-Info available |  |
|  | – meet requirements of 6.7 for creepage and – clearances between accessible parts and – hazardous live parts, if protection is provided by – limited access |  |  |
| 6.4.3 | basic insulation | (see Forms A.15 and A.16) |  |
|  | – meet clearance, creepage distance and solid insulation requirements of 6.7 |  |  |
| 6.4.4 | Impedance | (see Forms A.12 and A.15) |  |
|  | Impedance used as primary means of protection meets all the following requirements: |  | — |
|  | - limits current or voltage to level of 6.3.2 | (see Form A.6) |  |
|  | - rated for maximum working voltage and the amount of power it will dissipate | TBD-Info available |  |
|  | - clearance, creepage distance between terminations of the impedance meet requirements of basic insulation of 6.7 | (see Form A.15) |  |
| 6.5 | Additional means of protection in case of single fault condition |  |  |
| 6.5.1 | General |  |  |
|  | accessible parts are prevented from becoming hazardous live by the primary means of protection and supplemented by one of: |  | — |
|  | - protective bonding (see 6.5.2) |  |  |
|  | - supplementary insulation (see 6.5.3) |  |  |
|  | - automatic disconnection of the supply (see 6.5.5) | Fuse and grounding provide primary and supplementary protection. | P |
|  | - current- or voltage-limiting device (see 6.5.6) |  |  |
|  | Alternatively one of the single means of protection is used: |  | — |
|  | - reinforced insulation (see 6.5.3) |  |  |
|  | - protective impedance (see 6.5.4) |  |  |
| 6.5.2 | protective bonding | (see Forms A.7, A.8, A.9, A.10 or A.11) |  |
| 6.5.2.1 | General |  |  |
|  | accessible conductive parts, may become hazardous live in single fault condition: | TBD-Info available | — |
|  | Bonded to the protective conductor terminal; or | TBD-Info available |  |
|  | Separated by conductive screen or barrier bonded to protective conductor terminal | TBD-Info available |  |
| 6.5.2.2 | Integrity of protective bonding |  | — |
|  | - protective bonding consists of directly connected structural parts or discrete conductors or both; and withstands thermal and dynamic stresses | TBD-Info available |  |
|  | - Soldered connections: | TBD-Info available | — |
|  | Independently secured against loosening | TBD-Info available |  |
|  | Not used for other purposes | TBD-Info available |  |
|  | - Screw connections are secured | TBD-Info available |  |
|  | - protective bonding not interrupted; or | TBD-Info not available |  |
|  | except as removable part that carries mains supply input connection to the whole equipment | TBD-Info available |  |
|  | - Any movable protective bonding connection specifically designed, and meets 6.5.2.4 |  |  |
|  | - No external metal braid of cables used (not regarded as protective bonding) | TBD-Info available |  |
|  | - If mains supply passes through: | TBD-Info available | — |
|  | Means provided for passing protective conductor; | TBD-Info available |  |
|  | Impedance meets 6.5.2.4 |  |  |
|  | - Protective conductors bare or insulated, if insulated, green/yellow |  | N/A |
|  | Exceptions: |  | — |
|  | 1) earthing braids; | Critical information on conductor color or insulation missing. | N/A |
|  | 2) internal protective conductors etc.; | Critical information on conductor color or insulation missing. | N/A |
|  | Green/yellow not used for other purposes | No evidence of green/yellow used incorrectly found. | P |
|  | terminal suitable for connection of a protective conductor, and meets 6.5.2.3 |  |  |
| 6.5.2.3 | ## protective conductor terminal |  | — |
|  | - Contact surfaces are metal | TBD-Info available |  |
|  | - Appliance inlet used | TBD-Info available |  |
|  | - For rewirable cords and permanently connected equipment, protective conductor terminal is close to mains supply terminals | TBD-Info available |  |
|  | - If no mains supply is required, any protective conductor terminal: | TBD-Info available | — |
|  | Is near terminals of circuit for which protective earthing is necessary | TBD-Info available |  |
|  | External if other terminals external | TBD-Info available |  |
|  | - Equivalent current-carrying capacity to mains supply terminals | (see Form A.7) |  |
|  | - If plug-in, makes first and breaks last | TBD-Info available |  |
|  | - If also used for other bonding purposes, protective conductor: | TBD-Info available | — |
|  | Applied first; | TBD-Info not available |  |
|  | Secured independently; | TBD-Info available |  |
|  | Unlikely to be removed by servicing | TBD-Info available |  |
|  | - protective conductor of measuring circuit: | TBD-Info available | — |
|  | - Current rating equivalent to measuring circuit terminal; | TBD-Info available |  |
|  | - protective bonding: not interrupted by any switch or interrupting device | TBD-Info available |  |
|  | - functional earth terminals allow independent connection | TBD-Info available |  |
|  | - If a binding screw used for Protective conductor terminal: | TBD-Info available | — |
|  | Suitable size for bond wire | TBD-Info available |  |
|  | Not smaller than M 4 | TBD-Info available |  |
|  | At least 3 turns of screw engaged | TBD-Info available |  |
|  | Passes tightening torque test | (see Form A.8) |  |
|  | - Contact pressure not capable being reduced by deformation of materials |  |  |
| 6.5.2.4 | Impedance of protective bonding of plug-connected equipment | (see Form A.9) |  |
|  | Impedance between protective conductor terminal and each accessible part where protective bonding is specified, is: |  | — |
|  | – less than 0,1 Ohm; or | TBD-Info available |  |
|  | – less than 0,2 Ohm if equipment is provided with non-detachable cord | TBD-Info available |  |
| 6.5.2.5 | Impedance of protective bonding of permanently connected equipment | (see Form A.10) |  |
| 6.5.2.6 | Transformer protective bonding screen | (see Form A.11) |  |
|  | Transformer provided with screen for protective bonding: |  | — |
|  | screen bonding consists of directly connected structural parts or discrete conductors or both; and withstands thermal and dynamic stresses (see 6.5.2.2 a ) |  |  |
|  | screen bonding with soldered connection (see 6.5.2.2 b ) is: |  | — |
|  | – Independently secured against loosening | TBD-Info available |  |
|  | – Not used for other purposes | TBD-Info available |  |
| 6.5.3 | supplementary and reinforced insulation |  |  |
|  | Meet clearance, creepage distance and solid insulation requirements of 6.7 |  |  |
| 6.5.4 | protective impedance | (see Form A.12) |  |
|  | Limits current or voltage to level of 6.3.1 in normal and to level of 6.3.2 in single fault condition |  |  |
|  | clearance, creepage distance between terminations of the impedance meet requirements of double or reinforced insulation of 6.7 | (see Form A.15) |  |
|  | The protective impedance consists of one or more of the following: | (see TABLE 1.A and Form A.12) | — |
|  | - appropriate single component suitable for safety and reliability for protection, it is: | TBD-Info available | — |
|  | 1) rated twice the maximum working voltage | TBD-Info available |  |
|  | 2) resistor rated for twice the power dissipation for maximum working voltage | TBD-Info available |  |
|  | - combination of components | TBD-Info available |  |
|  | Single electronic device not used as protective impedance | TBD-Info available |  |
| 6.5.5 | Automatic disconnection of the supply |  | P |
|  | - rated to disconnect the load within time specified in Figure 2 | Fuse present, but disconnection time rating not shown. | N/A |
|  | - rated for the maximum load conditions of the equipment | Automatic disconnection device rating not documented in input files. | P |
| 6.5.6 | Current- or voltage-limiting devices | (see Form A.13) |  |
|  | Device complies with all of: |  | — |
|  | - rated to limit the current or voltage to the level of 6.3.2 | (see Form A.6) |  |
|  | - rated for the maximum working voltage; and | TBD-Info available |  |
|  | rated for the maximum operational current if applicable | TBD-Info available |  |
|  | - clearance, creepage distance between terminations of the impedance meet requirements of supplementary insulation of 6.7 | (see Forms A.14 and A.15) |  |
| 6.6 | Connections to external circuits |  |  |
| 6.6.1 | General |  |  |
|  | Connections do not cause accessible parts of the following to become hazardous live in normal condition or single fault condition: | TBD-Info available | — |
|  | – the external circuits | TBD-Info available |  |
|  | – the equipment | TBD-Info available |  |
|  | Protection achieved by separation of circuits; or | TBD-Info available |  |
|  | short circuit of separation does not cause a hazard | TBD-Info available |  |
|  | Instructions or markings for each terminal include: |  | — |
|  | - rated conditions for terminal | TBD-Info available |  |
|  | - Required rating of external circuit insulation | TBD-Info available |  |
| 6.6.2 | terminals for external circuits |  |  |
|  | terminals which receive a charge from an internal capacitor are not hazardous live after 10 s of interrupting supply connection | (see Form A.5) |  |
| 6.6.3 | Circuits with terminals which are hazardous live |  |  |
|  | These circuits are: |  | — |
|  | Not connected to accessible conductive parts; or | TBD-Info available |  |
|  | Connected to accessible conductive parts, but are not mains circuits and have one terminal contact at earth potential | TBD-Info available |  |
|  | No accessible conductive parts are hazardous live | TBD-Info available |  |
| 6.6.4 | Terminals for stranded conductors |  |  |
|  | No risk of accidental contact because: | TBD-Info available | — |
|  | – Located or shielded | TBD-Info available |  |
|  | – Self-evident or marked whether or not connected to accessible conductive parts | TBD-Info available |  |
|  | Complies as applicable: |  | — |
|  | - Manufacturer’s specified maximum length of removed insulation, or | TBD-Info available |  |
|  | - 8 mm length of insulation removed | TBD-Info available |  |
| 6.7 | Insulation requirements | (see Form A.14) |  |
| 6.7.1 | The nature of insulation |  |  |
| 6.7.1.1 | General |  |  |
|  | Insulation between accessible parts or between separate circuits consist of clearances, creepage distances and solid insulation if provided as protection against a hazard | TBD-Info available |  |
| 6.7.1.2 | clearances |  |  |
|  | Required clearances reflecting factors of 6.7.1.1 | (see Forms A.14 and A.15) |  |
|  | Equipment rated for operating altitude greater than 2000 m correction factor of Table 3 of 61010-1 applied | TBD-Info available |  |
| 6.7.1.3 | creepage distances |  |  |
|  | Required creepage distances reflecting factors of 6.7.1.1 a) to d) | (see Forms A.14 and A.15) |  |
|  | CTI material group reflected by requirements | TBD-Info available |  |
|  | CTI test performed | TBD-Info available |  |
| 6.7.1.4 | Solid insulation |  |  |
|  | Required solid insulation reflecting factors of 6.7.1.1 a) to d) | (see Forms A.14 and A.15) |  |
| 6.7.1.5 | Requirements for insulation according to type of circuit | (see Forms A.14 and A.15) |  |
|  | - 6.7.2 mains circuits of overvoltage category II up to nominal supply voltage of 300 V | TBD-Info available |  |
|  | - 6.7.3 secondary circuits separated from circuits defined in a) by transformer | TBD-Info available |  |
|  | - K.1 mains circuits of overvoltage category III and IV or overvoltage category II over 300 V | TBD-Info available |  |
|  | - K.2 secondary circuits separated from circuits defined in c) by transformer | TBD-Info available |  |
|  | - K.3 circuits having one or more of: |  | — |
|  | 1) maximum transient overvoltage is limited to known level below the level of mains circuit | TBD-Info available |  |
|  | 2) maximum transient overvoltage above the level of mains circuit | TBD-Info available |  |
|  | 3) Working voltage is the sum of more than one circuit or a mixed voltage | TBD-Info available |  |
|  | 4) Working voltage includes recurring peak voltage, may include non-sinusoidal or non-periodic waveform | TBD-Info available |  |
|  | 5) Working voltage with a frequency above 30 kHz | TBD-Info available |  |
| 6.7.2 | Insulation for mains circuits of overvoltage category II with a nominal supply voltage up to 300 V |  |  |
| 6.7.2.1 | clearances and creepage distances | (see Forms A.14 and A.15) | — |
|  | Values for mains circuits of Table 4 are met | TBD-Info available |  |
|  | Coatings to achieve reduction to pollution degree 1 comply with requirements of Annex H | TBD-Info available |  |
| 6.7.2.2 | Solid insulation |  |  |
| 6.7.2.2.1 | General |  |  |
|  | Withstands electrical and mechanical stresses in normal use and all rated environmental conditions of 1.4 | TBD-Info available |  |
|  | Equipment passed voltage tests of 6.8.3 with values of Table 5 | (see Form A.18) |  |
|  | Complies as applicable: |  | — |
|  | - enclosure or protective barrier of Clause 8 | TBD-Info available |  |
|  | - moulded and potted parts requirements of 6.7.2.2.2 |  |  |
|  | - inner layers of printed wiring boards requirements of 6.7.2.2.3 |  |  |
|  | - thin-film insulation requirements of 6.7.2.2.4 |  |  |
| 6.7.2.2.2 | Moulded and potted parts |  | — |
|  | Conductors between same two layers are separated by at least 0,4 mm after moulding is completed | TBD-Info available |  |
| 6.7.2.2.3 | Inner insulating layers of printed wiring boards |  | — |
|  | Separated by at least 0,4 mm between same two layers | TBD-Info available |  |
|  | reinforced insulation has adequate electric strength; one of following methods used: | TBD-Info available | — |
|  | - thickness of insulation is at least 0,4 mm | TBD-Info available |  |
|  | - insulation is assembled of minimum two separate layers, each rated for test voltage of Table 5 for basic insulation | TBD-Info available |  |
|  | - insulation is assembled of minimum two separate layers, where the combination is rated for test voltage of Table 5 for reinforced insulation | TBD-Info available |  |
| 6.7.2.2.4 | Thin-film insulation |  | — |
|  | Conductors between same two layers are separated by applicable clearances and creepage distance of 6.7.2.1 |  |  |
|  | reinforced insulation have adequate electric strength; one of the following methods used: | TBD-Info available | — |
|  | - thickness through the insulation at least 0,4 mm | TBD-Info available |  |
|  | - insulation is assembled of min. two separate layers, each rated for test voltage of Table 5 for basic insulation | (see Form A.18) |  |
|  | - insulation is assembled of min. three separate layers, where the combination of two layers passed voltage tests of 6.8.3 with values of Table 5 for reinforced insulation | (see Form A.18) |  |
| 6.7.3 | Insulation for secondary circuits derived from mains circuits of overvoltage category II up to 300 V | TBD-Info available |  |
| 6.7.3.1 | General |  |  |
|  | Secondary circuits where separation from mains circuits is achieved by a transformer providing: | TBD-Info available | — |
|  | – reinforced insulation | TBD-Info available |  |
|  | – double insulation | TBD-Info available |  |
|  | – screen connected to the protective conductor terminal | TBD-Info available |  |
| 6.7.3.2 | clearances | (see Forms A.14 and A.15) |  |
|  | - meet the values of Table 6 for basic insulation and supplementary insulation; or | TBD-Info available |  |
|  | twice the values of Table 6 for reinforced insulation; or | TBD-Info available |  |
|  | - pass the voltage tests of 6.8 with values of Table 6; | (see Form A.18) |  |
|  | with following adjustments: |  | — |
|  | 1) values for reinforced insulation are 1,6 times the values for basic insulation | TBD-Info available |  |
|  | 2) if operating altitude is greater than 2000 m values of clearances multiplied with factor of Table 3 | TBD-Info available |  |
|  | 3) minimum clearance is 0,2 mm for pollution degree 2 and 0,8 mm for pollution degree 3 | TBD-Info available |  |
| 6.7.3.3 | creepage distances | (see Forms A.14 and A.15) |  |
|  | Based on working voltage meets the values of Table 7 for basic and supplementary insulation | TBD-Info available |  |
|  | Values for reinforced insulation are twice the values of basic insulation | TBD-Info available |  |
|  | Coatings to achieve reduction to pollution degree 1 comply with requirements of Annex H | TBD-Info available |  |
| 6.7.3.4 | Solid insulation |  |  |
| 6.7.3.4.1 | General |  |  |
|  | Withstands electrical and mechanical stresses in normal use and all rated environmental conditions of 1.4 | TBD-Info available | — |
|  | - Equipment passed voltage test of 6.8.3.1 for 5 s with values of Table 6 for basic and supplementary insulation | (see Form A.18) |  |
|  | values for reinforced insulation are 1,6 times the values of basic insulation | TBD-Info available |  |
|  | - if working voltage exceeds 300 V, equipment passed voltage test of 6.8.3.1 for 1 min with a test voltage of 1,5 times working voltage for basic or supplementary insulation | (see Form A.18) |  |
|  | value for reinforced insulation are twice the working voltage | TBD-Info available |  |
|  | Complies as applicable: |  | — |
|  | 1) enclosure or protective barrier of Clause 8 | TBD-Info available |  |
|  | 2) moulded and potted parts requirements of 6.7.3.4.2 |  |  |
|  | 3) inner layers of printed wiring boards requirements of 6.7.3.4.3 |  |  |
|  | 4) thin-film insulation requirements of 6.7.3.4.4 |  |  |
| 6.7.3.4.2 | Moulded and potted parts |  | — |
|  | Conductors between same two layers are separated by applicable distances of Table 8 | TBD-Info available |  |
| 6.7.3.4.3 | Inner insulation layers of printed wiring boards |  | — |
|  | Separated by at least the applicable distances of Table 8 between same two layers | TBD-Info available |  |
|  | reinforced insulation have adequate electric strength; one of following methods used: | TBD-Info available | — |
|  | - thickness at least applicable distance of Table 8 | TBD-Info available |  |
|  | - insulation is assembled of minimum two separate layers, each rated for test voltage of Table 6 for basic insulation | (see Form A.18) |  |
|  | - insulation is assembled of min. two separate layers, where the combination is rated for 1,6 times the test voltage of Table 6 | (see Form A.18) |  |
| 6.7.3.4.4 | Thin-film insulation |  | — |
|  | Conductors between same two layers are separated by applicable clearances and creepage distance of 6.7.3.2 and 6.7.3.3 |  |  |
|  | reinforced insulation have adequate electric strength; one of following methods used: | TBD-Info available | — |
|  | - thickness at least applicable distance of Table 8 | TBD-Info available |  |
|  | - insulation is assembled of min. two separate layers, each rated for test voltage of Table 6 for basic insulation | (see Form A.18) |  |
|  | - insulation is assembled of min. three separate layers, where the combination of two layers passed voltage tests with 1,6 time values of Table 6: | (see Form A.18) | — |
|  | a.c. test of 6.8.3.1; or |  |  |
|  | d.c. test of 6.8.3.2 for circuits stressed only by d.c. voltages |  |  |
| 6.8 | Procedure for voltage tests | (see Forms A.14 and A.18) |  |
| 6.9 | Constructional requirements for protection against electric shock |  |  |
| 6.9.1 | General |  |  |
|  | If a failure could cause a hazard: | TBD-Info available | — |
|  | - security of wiring connections | TBD-Info available |  |
|  | - screws securing removable covers | TBD-Info available |  |
|  | - accidental loosening | TBD-Info available |  |
|  | - clearances and creepage distances not reduced below the values of basic insulation by loosening of parts or wires | TBD-Info available |  |
| 6.9.2 | Insulating materials |  |  |
|  | Material not to be used for safety relevant insulation: |  | — |
|  | - easily damaged materials not used |  |  |
|  | - non-impregnated hygroscopic materials not used | TBD-Info not available |  |
| 6.9.3 | Colour coding |  |  |
|  | Green-and-yellow insulation shall not be used except: |  | — |
|  | - protective earth conductors; | No evidence of green-yellow insulation misuse found. | P |
|  | - protective bonding conductors; | Green-yellow used only for protective bonding conductors. | N/A |
|  | - potential equalization conductors; | No evidence of green-yellow insulation misuse found. | N/A |
|  | - functional earth conductors | No evidence of green-yellow insulation misuse found. | P |
| 6.10 | Connection to mains supply source and connections between parts of equipment |  |  |
| 6.10.1 | mains supply cords |  |  |
|  | rated for maximum equipment current (see 5.1.3 c) | TBD-Info available |  |
|  | Cable complies with IEC 60227 or IEC 60245 | Cable standard compliance evidence not available. | N/A |
|  | Heat-resistant if likely to contact hot parts | TBD-Info available |  |
|  | Temperature rating (cord and inlet) : | TBD-Info available | — |
|  | Green/yellow used only for connection to protective conductor terminals | TBD-Info available |  |
|  | Detachable cords with IEC 60320 mains connectors: | No evidence of IEC 60320 connectors provided. | — |
|  | Conform to IEC 60799; or | Critical mains cord details not available in documentation. |  |
|  | Have the current rating of the mains connector | TBD-Info available |  |
| 6.10.2 | Fitting of non-detachable mains supply cords |  |  |
| 6.10.2.1 | Cord entry |  | — |
|  | - inlet or bushing with a smoothly rounded opening; or | Critical information on cord fitting not available. | N/A |
|  | - insulated cord guard protruding >5 D (diameter) | Critical information on cord guard not available. | N/A |
| 6.10.2.2 | Cord anchorage |  | — |
|  | Protective earth conductor is the last to take the strain | TBD-Info available |  |
|  | - cord is not clamped by direct pressure from a screw | TBD-Info available |  |
|  | - knots are not used | TBD-Info available |  |
|  | - cannot push the cord into the equipment to cause a hazard | TBD-Info available |  |
|  | - no failure of cord insulation in anchorage with metal parts | TBD-Info available |  |
|  | - not to be loosened without a tool | TBD-Info available |  |
|  | - cord replacement does not cause a hazard and method of strain relief is clear | TBD-Info available |  |
|  | Push-pull and or torque test | (see Form A.19) |  |
| 6.10.3 | Plugs and connectors |  |  |
|  | mains supply plugs, connectors etc., conform with relevant specifications | TBD-Info available |  |
|  | If equipment supplied at voltages below 6.3.2.a) or from a sole source: |  | — |
|  | Plugs of supply cords do not fit mains sockets above rated supply voltage | TBD-Info available |  |
|  | Mains type plugs used only for connection to mains supply | TBD-Info available |  |
|  | Plug pins which receive a charge from an internal capacitor | (see Form A.5) |  |
|  | Accessory mains socket outlets: | TBD-Info available | — |
|  | - marking if accepts a standard mains supply plug (see 5.1.3e) | No accessory mains socket-outlets identified in documentation. | N/A |
|  | - input has a protective earth conductor if outlet has earth terminal contact | TBD-Info available |  |
| 6.11 | Disconnection from supply source |  |  |
| 6.11.1 | Disconnects all current-carrying conductors | TBD-Info available |  |
| 6.11.2 | Exceptions | TBD-Info available |  |
| 6.11.3 | Requirements according to type of equipment |  |  |
| 6.11.3.1 | permanently connected equipment and multi-phase equipment | Disconnection means not explicitly shown in documentation. |  |
|  | Employs switch or circuit-breaker | No switch or circuit-breaker evidence found. | N/A |
|  | If switch or circuit-breaker is not part of the equipment, documentation requires: |  | — |
|  | - switch or circuit-breaker to be included in building installation | Critical information not available in the documentation. | N/A |
|  | - suitable location easily reached | Critical information not available in the documentation. | N/A |
|  | - marking as disconnecting for the equipment | Critical marking information not available in documentation. | N/A |
| 6.11.3.2 | Single-phase cord-connected equipment |  |  |
|  | Equipment is provided with one of the following: | Power switch present; fuse shown before power supply. | — |
|  | - switch or circuit-breaker | Presence of power switch is indicated in TRF text. | P |
|  | - appliance coupler (disconnectable without tool) | Critical information on appliance coupler not available. | N/A |
|  | - separable plug (without locking device) | Critical information on plug type is not available. | P |
| 6.11.4 | Disconnecting devices |  |  |
| 6.11.4.1 | General |  | P |
|  | Disconnecting device part of equipment | Power switch is shown as part of equipment. | P |
|  | Electrically close to the supply | Disconnecting device is electrically close to supply. | P |
|  | Power-consuming components not electrically located between the supply source and the disconnecting device | Fuse is before disconnecting device; compliant. | P |
|  | Except electromagnetic interference suppression circuits permitted to be located on the supply side of the disconnecting device | All circuits are on load side of disconnect device. | P |
| 6.11.4.2 | Switches and circuit-breakers |  |  |
|  | When used as disconnection device: |  | — |
|  | Circuit breaker meets the relevant requirements IEC 60947-2 and is suitable for the application | Critical circuit breaker information not available in documentation. | N/A |
|  | Switch meets the relevant requirements IEC 60947-3 and is suitable for the application | Critical information on switch compliance is not available. |  |
|  | Marked to indicate function : | Critical marking information not available in documentation. | — |
|  | Not incorporated in mains cord | No disconnecting device found in mains cord evidence. | P |
|  | Does not interrupt protective earth conductor |  |  |
| 6.11.4.3 | Appliance couplers and plugs |  |  |
|  | Where an appliance coupler or separable plug is used as the disconnecting device (see 6.11.3.2): | Presence of power switch is confirmed in TRF text. | — |
|  | Readily identifiable and easily reached by the operator | Critical information on coupler accessibility not available. | P |
|  | Single-phase portable equipment cord length not more than 3 m | Cord length information not available in documentation. | N/A |
|  | protective earth conductor connected first and disconnected last | Critical information on coupler sequence not available. | P |
| 7 | PROTECTION AGAINST MECHANICAL HAZARDS |  |  |
| 7.1 | General |  |  |
|  | Equipment does not cause a mechanical hazard in normal nor in single fault condition | Mechanical hazard evidence not available in documentation. | N/A |
|  | Conformity is checked by 7.2 to 7.7 |  |  |
| 7.2 | Sharp edges |  | N/A |
|  | Easily-touched parts are smooth and rounded | Critical information on edges and surfaces is missing. | N/A |
|  | Do not cause injury during normal use and | Critical information on sharp edges is not available. | N/A |
|  | Do not cause injury during single fault condition | Critical information on sharp edges is not available. | N/A |
| 7.3 | Moving parts |  |  |
| 7.3.1 | General |  |  |
|  | hazards from moving parts limited to a tolerable level with the conditions specified in 7.3.2 and 7.3.5 |  |  |
|  | risk assessment in accordance with 7.3.3 carried out |  |  |
| 7.3.2 | Exceptions |  |  |
|  | Access to hazardous moving parts permitted under following circumstances: |  | — |
|  | - obviously intended to operate on parts or materials external of the equipment | Critical information on hazardous moving parts is missing. | N/A |
|  | inadvertent touching of moving parts minimized by equipment design (e .g. guards or handles) | Critical information on moving parts is not available. | N/A |
|  | - If operator access is unavoidable outside normal use following precautions have been taken: |  | — |
|  | 1) access requires tool | Critical information on hazardous moving parts is not available. | N/A |
|  | 2) statement about training in the instructions | Critical information not available in the documentation. | N/A |
|  | 3) warning markings on covers prohibiting access by untrained operators | Critical information on hazardous moving parts is not available. | N/A |
|  | or symbol 14 with full details in documentation | Critical information on hazardous moving parts is missing. | N/A |
| 7.3.3 | risk assessment for mechanical hazards to body parts |  |  |
|  | risk is reduced to a tolerable level by protective measures as specified in Table 12 | TBD-Info available |  |
|  | Minimum protective measures: |  | — |
|  | - Low level measures | TBD-Info available |  |
|  | - Moderate measures | TBD-Info available |  |
|  | - Stringent measures | TBD-Info not available |  |
| 7.3.4 | Limitation of force and pressure | (see Form A.20) |  |
|  | Following levels are met in normal and single fault condition: |  | — |
|  | Continuous contact pressure below 50 N / cm² with force below 150 N | TBD-Info available |  |
|  | Temporary force below 250 N for an area at least of 3 cm² for a maximum duration of 0,75 s | TBD-Info available |  |
| 7.3.5 | Gap limitations between moving parts | (see Form A.20) |  |
| 7.3.5.1 | Access normally allowed |  | — |
|  | If levels of 7.3.4 exceeded and a body part may be inserted minimum gap as specified in Table 13 assured in normal and in single fault condition |  |  |
| 7.3.5.2 | Access normally prevented |  | — |
|  | Maximum gap as specified in Table 14 assured in normal and in single fault condition | TBD-Info available |  |
| 7.4 | Stability |  |  |
|  | Equipment not secured to building structure is physical stable | Critical information on physical stability is not available. | N/A |
|  | Stability maintained after opening of drawers etc. by automatic means, or | Critical information on stability provisions is not available. | N/A |
|  | warning marking requires the application of means | Critical information on stability markings not available. | N/A |
|  | Compliance checked by following tests as applicable: | (see Form A.20A) | — |
|  | - 10° tilt test for other than handheld equipment | TBD-Info available |  |
|  | - multi-directional force test for equipment exceeds height of 1 m and mass of 25 kg | TBD-Info available |  |
|  | - downward force test for floor-standing equipment | TBD-Info available |  |
|  | - overload test with 4 times maximum load for castor or support foot that supports greatest load, or | TBD-Info available |  |
|  | - castor or support foot that supports greatest load removed from equipment | TBD-Info available |  |
| 7.5 | Provisions for lifting and carrying |  |  |
| 7.5.1 | General |  | N/A |
|  | Equipment more than 18 kg : | Critical information on weight and provisions not available. | N/A |
|  | Has means for lifting or carrying; or | Critical information on lifting means not available. | N/A |
|  | Directions are given in documentation | Directions for lifting and carrying not found in documentation. | N/A |
| 7.5.2 | Handles and grips |  | N/A |
|  | Handles or grips withstand four times weight | Critical information on handles/grips strength is not available. | N/A |
| 7.5.3 | Lifting devices and supporting parts |  |  |
|  | rated for maximum load; or | TBD-Info available |  |
|  | Tested with four times maximum static load | TBD-Info available |  |
| 7.6 | Wall mounting |  |  |
|  | Mounting brackets withstand four times weight | (see Form A.20B) |  |
|  | One fastener removed and test repeated with two times weight | (see Form A.20B) |  |
| 7.7 | Expelled parts |  | N/A |
|  | Equipment contains or limits the energy | Critical information on expelled parts is not available. | N/A |
|  | Protection not removable without the aid of a tool | Critical information on expelled parts protection is missing. | N/A |
| 8 | RESISTANCE TO MECHANICAL STRESSES |  |  |
| 8.1 | General |  |  |
|  | Equipment does not cause a hazard when subjected to mechanical stresses in normal use | Critical mechanical stress test evidence not available in documentation. | N/A |
|  | Normal protection level is 5 J | Critical information on 5 J impact test not available. | N/A |
|  | Levels below 5 J but not less than 1 J are acceptable if all of the following criteria are met: |  | — |
|  | - Lower level justified by risk assessment of manufacturer | Risk assessment justification for <5 J impact not found. | N/A |
|  | - Equipment installed in its intended application is not easily touched | Critical information on touch accessibility is not available. | N/A |
|  | - Only occasional access during normal use | Critical information on access frequency is not available. | N/A |
|  | - IK code in accordance to IEC 62262 marked or symbol 14 used with full information in the documentation |  |  |
|  | for non-metallic enclosures rated below 2 °C ambient temperature value chosen for minimum rated temperature |  |  |
|  | impact energies between IK values, the IK code marked for nearest lower value |  |  |
|  | Conformity is checked by performing following tests: | (see Form A.16) | — |
|  | - Static test of 8.2.1 |  |  |
|  | - Impact test of 8.2.2 with 5 J except for hand-held equipment |  |  |
|  | if specified impact energy is not 5 J alternate method of IEC 62262 used |  |  |
|  | - Drop test of 8.3.1 or 8.3.2 except for fixed equipment and equipment with mass over 100 kg |  |  |
|  | Equipment rated with an impact rating of IK 08 that obviously meets the criteria | TBD-Info available |  |
|  | After the tests inspection with following results: |  | — |
|  | – hazardous live parts above the limits of 6.3.2 not accessible |  |  |
|  | – insulation pass the voltage tests of 6.8 | (see Form A.30) |  |
|  | - No leaks of corrosive and harmful substances | TBD-Info available |  |
|  | - enclosure shows no cracks resulting in a hazard | TBD-Info available |  |
|  | - clearances not less than their permitted values | TBD-Info available |  |
|  | - Insulation of internal wiring remains undamaged | TBD-Info available |  |
|  | - protective barriers not damaged or loosened | TBD-Info available |  |
|  | - No moving parts exposed, except permitted by 7.3 | TBD-Info available |  |
|  | - No damage which could cause spread of fire | TBD-Info available |  |
| 8.2 | Enclosure rigidity test |  |  |
| 8.2.1 | Static test | (see Form A.21A) |  |
|  | – 30 N with 12 mm rod applied to each part of enclosure | TBD-Info available |  |
|  | – in case of doubt test conducted at maximum rated ambient temperature | TBD-Info available |  |
| 8.2.2 | Impact test | (see Form A.21A) |  |
|  | Impact applied to any part of enclosure causing a hazard if damaged | TBD-Info available |  |
|  | Impact energy level and corresponding IK code : | TBD-Info available | — |
|  | Non-metallic enclosures cooled to minimum rated ambient temperature if below 2 °C | TBD-Info available |  |
| 8.3 | Drop test | (see Form A.21B) |  |
| 8.3.1 | Other than hand-held and direct-plug-in equipment | TBD-Info available |  |
|  | Tests conducted with a drop height or angle of : | TBD-Info available | — |
| 8.3.2 | hand-held and direct-plug-in equipment | TBD-Info available |  |
|  | Non-metallic enclosures cooled to minimum rated ambient temperature if below 2 °C | TBD-Info available |  |
|  | Drop test conducted with an height of 1 m | TBD-Info available |  |
| 9 | PROTECTION AGAINST THE SPREAD OF FIRE |  |  |
| 9.1 | General |  |  |
|  | No spread of fire in normal and single fault condition | Critical fire spread evidence not available in documentation. | N/A |
|  | mains supplied equipment meets requirements of 9.6 additionally |  |  |
|  | Conformity is checked by minimum one or a combination of the following (see Figure 11): | (see Form A.22) | — |
|  | - Single Fault test of 4.4; or | (see Form A.1) |  |
|  | - Application of 9.2 (eliminating or reducing the sources of ignition); or |  |  |
|  | - Application of 9.3 (containment of fire within the equipment) |  |  |
| 9.2 | Eliminating or reducing the sources of ignition within the equipment |  |  |
|  | - 1) Limited-energy circuit (see 9.4); or |  |  |
|  | 2) basic insulation provided for parts of different potential; or | (see Forms A.14 and A.18) |  |
|  | Bridging the insulation does not cause ignition | (see Form A.1) |  |
|  | - Surface temperature of liquids and parts (see 9.5) |  |  |
|  | - No ignition in circuits designed to produce heat | (see Form A.1) |  |
| 9.3 | Containment of the fire within the equipment, should it occur |  |  |
| 9.3.1 | General |  |  |
|  | Spread of fire outside equipment reduced to a tolerable level if: |  | — |
|  | - Energizing of the equipment is controlled by an operator held switch | Critical fire containment evidence not available in documentation. | P |
|  | - enclosure is conform with constructional requirements of 9.3.2; and |  |  |
|  | Requirements of 9.5 are met |  |  |
| 9.3.2 | Constructional requirements |  |  |
|  | - Connectors and insulating material have flammability classification V-2 or better | (see TABLE 1.A or Form A.23) |  |
|  | - Insulated wires and cables are flame retardant (VW-1 or equivalent) | (see TABLE 1.A or Form A.23) |  |
|  | - enclosure meets following requirements: | (see Form A.22) | — |
|  | - Bottom and sides in arc of 5 ° (see Figure 13) to non-limited circuits (9.4) meets: | TBD-Info available | — |
|  | - no openings; or | TBD-Info available |  |
|  | - perforated as specified in Table 16; or | TBD-Info available |  |
|  | - metal screen with a mesh; or | TBD-Info available |  |
|  | - baffles as specified in Figure 12 | TBD-Info available |  |
|  | - Material of enclosure and any baffle or flame barrier is made of: | TBD-Info available | — |
|  | Metal (except magnesium); or | TBD-Info available |  |
|  | Non-metallic materials have flammability classification V-1 or better | (see TABLE 1.A or Form A.22) |  |
|  | - Enclosure and any baffle or flame barrier have adequate rigidity | TBD-Info available |  |
| 9.4 | Limited-energy circuit | (see Form A.24) |  |
|  | - Potential not more than 30 r.m.s. and 42,4 V peak, or 60 V d.c. | TBD-Info available |  |
|  | - Current limited by one of following means: | TBD-Info available | — |
|  | - Inherently or by impedance (see Table 17); or | TBD-Info available |  |
|  | - Overcurrent protective device (see Table 18); or | TBD-Info available |  |
|  | - A regulating network limits also in single fault condition (see Table 17) | TBD-Info available |  |
|  | - Is separated by at least basic insulation | TBD-Info available |  |
|  | Fuse or a nonadjustable electromechanical device is used | TBD-Info available |  |
| 9.5 | Requirements for equipment containing or using flammable liquids |  |  |
|  | Flammable liquids contained in or specified for use with equipment do not cause spread of fire | (see Form A.25) |  |
|  | risk is reduced to a tolerable level: |  | — |
|  | - The temperature of surface or parts in contact with flammable liquids is 25 °C below fire point | TBD-Info available |  |
|  | - The quantity of liquid is limited | TBD-Info available |  |
|  | - Flames are contained within the equipment | TBD-Info available |  |
|  | Detailed instructions for risk-reduction provided | TBD-Info available |  |
| 9.6 | Overcurrent protection |  |  |
| 9.6.1 | General |  |  |
|  | mains supplied equipment protected | Fuse present at AC input for overcurrent protection. | N/A |
|  | basic insulation between mains parts of opposite polarity provided | (see Forms A.14 and A.15) |  |
|  | Overcurrent protection devices not fitted in the protective conductor | TBD-Info not available |  |
|  | Fuses or single-pole circuit-breakers not fitted in neutral (multi-phase equipment) | TBD-Info available |  |
| 9.6.2 | permanently connected equipment |  |  |
|  | Overcurrent protection device: | Fuse present at AC input for overcurrent protection. | — |
|  | Fitted within the equipment; or | Fuse shown in block diagram as overcurrent protection. | P |
|  | Specified in manufacturer's instructions | Overcurrent protection device not specified in instructions. | P |
| 9.6.3 | Other equipment |  | N/A |
|  | Protection within the equipment | Fuse present at AC input for overcurrent protection. | N/A |
| 10 | EQUIPMENT TEMPERATURE LIMITS AND RESISTANCE TO HEAT |  |  |
| 10.1 | Surface temperature limits for protection against burns |  |  |
|  | Easily touched surfaces within the limits in normal and in single fault condition: | (see Form A.26A) | — |
|  | – at an specified ambient temperature of 40 °C | TBD-Info available |  |
|  | – for equipment rated above 40 °C ambient temperature limits not exceeded raised by the difference to 40 °C | TBD-Info available |  |
|  | Heated surfaces necessary for functional reasons exceeding specified values: | TBD-Info not available | — |
|  | – Are recognizable as such by appearance or function; or | TBD-Info available |  |
|  | – Are marked with symbol 13 | TBD-Info available |  |
|  | – Guards are not removable without tool | TBD-Info available |  |
| 10.2 | Temperatures of windings |  |  |
|  | Limits not exceeded in: | (see Form A.26B) | — |
|  | normal condition | TBD-Info available |  |
|  | single fault condition | TBD-Info available |  |
| 10.3 | Other temperature measurements |  |  |
|  | Following measurements conducted if applicable: | (see Form A.26A) | — |
|  | - Value of 60 °C of field-wiring terminal box not exceeded | TBD-Info available |  |
|  | - Surface of flammable liquids and parts in contact with this liquids | TBD-Info available |  |
|  | - Surface of non-metallic enclosures | TBD-Info available |  |
|  | - Parts made of insulating material supporting parts connected to mains supply | TBD-Info available |  |
|  | - Terminals carrying a current more than 0,5 A | TBD-Info available |  |
| 10.4 | Conduct of temperature tests |  |  |
| 10.4.1 | General |  |  |
|  | Tests conducted under reference test conditions and manufacturer’s instructions | (see Form A.26A) |  |
|  | Tests alternatively conducted at the least favourable ambient temperature within the rated ambient temperature : | TBD-Info available | — |
| 10.4.2 | Temperature measurement of heating equipment |  |  |
|  | Tests conducted in test corner | (see Form A.26A) |  |
| 10.4.3 | Equipment intended for installation in a cabinet or wall |  |  |
|  | Equipment built in as specified in installation instructions | (see Form A.26A) |  |
| 10.5 | Resistance to heat |  |  |
| 10.5.1 | Integrity of clearance and creepage distances | (see Form A.16) |  |
| 10.5.2 | Non-metallic enclosures | (see Form A.27) |  |
|  | Within 10 min after treatment: | TBD-Info available | — |
|  | Equipment subjected to suitable stresses of 8.2 and 8.3 complying with criteria of 8.1 | TBD-Info available |  |
| 10.5.3 | Insulating material | TBD-Info available |  |
|  | - Parts supporting parts connected to mains supply | Critical material data for insulation not available. | N/A |
|  | - terminals carrying a current more than 0,5 A | Critical information on insulating material heat resistance missing. | N/A |
|  | Examination of material data; or | Critical material data for terminals not available. | N/A |
|  | in case of doubt: |  |  |
|  | - Ball pressure test; or | (see Form A.28) |  |
|  | - Vicat softening test of ISO 306 | (see Form A.29) |  |
| 11 | PROTECTION AGAINST HAZARDS FROM FLUIDS AND SOLID FOREIGN OBJECTS |  |  |
| 11.1 | General |  | P |
|  | Protection to operators and surrounding area provided by equipment | No evidence of fluid/solid ingress protection found. | N/A |
|  | All fluids specified by manufacturer considered | Liquid and gas fluids are specified and considered. | P |
| 11.2 | Cleaning | (see Form A.30) |  |
| 11.3 | Spillage | (see Form A.30) |  |
| 11.4 | Overflow | (see Form A.30) |  |
| 11.5 | Battery electrolyte | TBD-Info available |  |
|  | Battery electrolyte leakage presents no hazard | No evidence of battery electrolyte leakage hazard provided. | P |
| 11.6 | Equipment rated with a degree of ingress protection (IP code) | (see Form A.30) |  |
| 11.6.1 | General |  |  |
|  | Equipment marked with IP code : | IP code marking evidence not found in documentation or images. | — |
|  | Conditions specified in the documentation | IP code rating information not found in documentation. | N/A |
| 11.6.2 | Conditions for testing |  |  |
|  | Equipment in clean and new condition, all parts in place and mounted as specified by manufacturer | TBD-Info available |  |
|  | Complete equipment tested, or | TBD-Info available |  |
|  | representative parts tested | TBD-Info available |  |
|  | hand-held equipment and portable equipment placed in least favourable position of normal use | TBD-Info available |  |
|  | Other equipment positioned or installed as specified | TBD-Info available |  |
|  | terminals provided with protective cap or cover, are installed as specified by manufacturer | TBD-Info available |  |
|  | The equipment is operating (energized) during the treatment except: | TBD-Info available | — |
|  | - If manufacturer specifies degrees of protection for non-operating (de-energized) equipment, or | TBD-Info available |  |
|  | - Equipment is operating or non-operating during the treatment with does not affect the test results | TBD-Info available |  |
| 11.6.3 | Protection against solid foreign objects (including dust) |  |  |
|  | Applicable test of IEC 60529 for protection against solid foreign objects conducted |  |  |
|  | Additionally inspection of equipment resulted: |  | — |
|  | - No deposit on insulation parts that could lead to a hazard |  |  |
|  | - No created accumulations that have the potential to cause spread of fire |  |  |
| 11.6.4 | Protection against water |  |  |
|  | Applicable test of IEC 60529 for protection against water conducted |  |  |
|  | If any water has entered, safety is not impaired, inspection of equipment resulted: |  | — |
|  | - No deposit on insulation parts that could lead to a hazard |  |  |
|  | - Water has not reached hazardous live parts or windings which are not designed to operate when wet |  |  |
|  | - No accumulations near the end of cable nor enter the cable where it could cause a hazard |  |  |
|  | - No accumulations where it could lead to a hazard taking in consideration movement of the equipment |  |  |
| 11.7 | Fluid pressure and leakage |  |  |
| 11.7.1 | Maximum pressure : | (see Form A.31) | — |
|  | Maximum pressure of any part does not exceed Prated | TBD-Info not available |  |
| 11.7.2 | Leakage and rupture at high pressure |  |  |
|  | Fluid-containing parts checked by inspection or if a hazard could arise subjected to hydraulic test, if: | (see Form A.31) | — |
|  | - product of pressure and volume > 200 kPa·l; and | TBD-Info available |  |
|  | - pressure > 50 kPa | TBD-Info not available |  |
|  | Safety evidence established by calculation in acc. to national authorities (e.g. Pressure Equipment Directive 2014/68/EU) | TBD-Info available |  |
|  | Parts of refrigerating systems meets pressure-related requirements of EN 378-2 or IEC 60335-2-89 as applicable |  |  |
| 11.7.3 | Leakage from low-pressure parts | (see Form A.32) |  |
| 11.7.4 | Overpressure safety device | TBD-Info available |  |
|  | Does not operate in normal use | TBD-Info available |  |
|  | - Connected as close as possible to parts intended to be protected | TBD-Info available |  |
|  | - Easy access for inspection, maintenance and repair | TBD-Info available |  |
|  | - Adjustment only with tool | TBD-Info available |  |
|  | - No discharge towards person | TBD-Info available |  |
|  | - No hazard from deposit of discharged material | TBD-Info available |  |
|  | - Adequate discharge capacity | TBD-Info available |  |
|  | No shut-off valve between overpressure safety device and protected parts | TBD-Info available |  |
| 12 | Protection against radiation, including laser sources, and against sonic and ultrasonic pressure |  |  |
| 12.1 | General |  |  |
|  | Equipment provides protection | TBD-Info available |  |
| 12.2 | Equipment producing ionizing radiation |  |  |
| 12.2.1 | Ionizing radiation | (see Form A.33) |  |
| 12.2.1.1 | General |  |  |
|  | Equipment meets the following requirements: |  | — |
|  | - if intended to emit radiation meets requirements of 12.2.1.2; or | TBD-Info available |  |
|  | tested, classified and marked in accordance to IEC 62598 |  |  |
|  | - if only emits stray radiation meets requirements of 12.2.1.3 | TBD-Info available |  |
| 12.2.1.2 | Equipment intended to emit radiation |  | — |
|  | Effective dose rate of radiation measured : | TBD-Info available | — |
|  | If dose rate exceeds 5 µSv/h marked with the following: | TBD-Info available | — |
|  | - symbol 17 (ISO 361) | TBD-Info available |  |
|  | - abbreviations of the radionuclides : | TBD-Info available | — |
|  | - with maximum dose at 1 m; or : | TBD-Info available | — |
|  | with dose rate value between 1 µSv/h and 5 µSv/h in m : | TBD-Info available | — |
| 12.2.1.3 | Equipment not intended to emit radiation | (see Form A.34) | — |
|  | Limit for unintended stray radiation of 1 µSv/h at any easily reached point kept : | TBD-Info available |  |
| 12.2.2 | Accelerated electrons | TBD-Info available |  |
|  | Compartments opened only by the use of a tool | TBD-Info available |  |
| 12.3 | Optical radiation |  |  |
|  | No unintentional hazardous escape of optical radiation as ultraviolet, visible or infrared radiation, including light emitting diodes: |  | — |
|  | – Checked by inspection; and |  |  |
|  | – Radiation sources assessed in acc. to the requirements of IEC 62471, except for sources considered to be safe (Table 22) or conditionally safe (Table 23). |  |  |
|  | – Lamp and lamp systems assessed to Risk Groups 1, 2, or 3 of IEC 62471 are labelled in acc. to IEC 62471-2 |  |  |
|  | – If labelling impractical, lamp or lamp systems marked with symbol 14 | Critical information on optical radiation is not available. | P |
|  | – Protective measures, restrictions on use, and operating instructions that may be necessary are provided, including the applicable conditions of use of Table 23. | Critical information on optical radiation hazards not available. | P |
| 12.4 | Microwave radiation |  |  |
|  | Power density does not exceed 10 W/m2 : | TBD-Info available |  |
| 12.5 | Sonic and ultrasonic pressure |  |  |
| 12.5.1 | Sound level | (see Form A.35) |  |
|  | No hazardous sound emission | TBD-Info available |  |
|  | Maximum sound pressure level measured and calculated for maximum sound power level as specified in ISO 3746 or ISO 9614-1 |  |  |
|  | Instruction describes measures for protection | TBD-Info available |  |
| 12.5.2 | Ultrasonic pressure | (see Form A.36) |  |
|  | Equipment not intended to emit ultrasound does not exceed limit of 110 dB between 20 kHz and 100 kHz | TBD-Info available |  |
|  | Equipment intended to emit ultrasound: | TBD-Info available |  |
|  | Outside useful beam does not exceed limit of 110 dB between 20 kHz and 100 kHz | TBD-Info available |  |
|  | If inside useful beam above values exceeded: | TBD-Info available | — |
|  | Marked with Symbol 14 of Table 1 | TBD-Info available |  |
|  | and following information in the documentation: | TBD-Info available | — |
|  | - dimensions of useful beam | TBD-Info available |  |
|  | - area where ultrasonic pressure exceed 110 dB | TBD-Info available |  |
|  | - maximum sound pressure inside beam area | TBD-Info available |  |
| 12.6 | Laser sources |  |  |
|  | Equipment meets requirements of IEC 60825-1 |  |  |
| 13 | PROTECTION AGAINST LIBERATED GASES AND SUBSTANCES, EXPLOSION AND IMPLOSION |  |  |
| 13.1 | Poisonous and injurious gases and substances |  |  |
|  | No hazardous substances liberated in normal condition and in single fault condition | No evidence of hazardous gases or substances found. | P |
|  | If potentially-hazardous substances are liberated: |  | — |
|  | Operator is not directly exposed to a quantity of the substance that could cause harm |  |  |
|  | Requirements to discharge of hazardous substances during normal operation in accordance to manufacturer’s instructions not considered as liberation |  |  |
|  | Attached data/test reports demonstrate conformity |  |  |
| 13.2 | Explosion and implosion |  |  |
| 13.2.1 | Components |  |  |
|  | Components liable to explode: |  | — |
|  | Pressure release device provided; or | Pressure relief safety valves are provided for pneumatic system. | P |
|  | Apparatus incorporates operator protection (see also 7.7) | No evidence of explosion-prone components or operator protection. | N/A |
|  | Pressure release device: |  | — |
|  | Discharge without danger | Pressure relief safety valves present for risk components. | N/A |
|  | Cannot be obstructed | No explosion/implosion protection components identified in documentation. | N/A |
| 13.2.2 | Batteries and battery charging | (see Form A.37) |  |
|  | If explosion or fire hazard could occur: |  | — |
|  | Protection incorporated in the equipment; or | TBD-Info available |  |
|  | Instructions specify batteries with built-in protection | TBD-Info available |  |
|  | In case of wrong type of battery used: | TBD-Info available | — |
|  | No hazard; or | TBD-Info available |  |
|  | Warning by marking and within instructions | TBD-Info available |  |
|  | Equipment with means to charge rechargeable batteries: | TBD-Info available | — |
|  | Warning against the charging of non-rechargeable batteries; and | TBD-Info available |  |
|  | Type of rechargeable battery indicated; or | TBD-Info available |  |
|  | Symbol 14 used | TBD-Info available |  |
|  | Battery compartment design | TBD-Info available |  |
|  | Single component failure | TBD-Info available |  |
|  | Polarity reversal test | TBD-Info available |  |
| 13.2.3 | Implosion of cathode ray tubes |  |  |
|  | If maximum face dimensions > 160 mm : | Face dimensions information not available in documentation. | — |
|  | Intrinsically protected and correctly mounted; or | No intrinsically protected equipment identified in documentation. | P |
|  | enclosure provides protection: | Metal enclosure provides electric shock protection. | P |
|  | If non-intrinsically protected: |  | — |
|  | Screen not removable without tool | Critical information on screen removal not available. | N/A |
|  | If glass screen, not in contact with surface of tube | No glass screen or tube information available. | N/A |
| 14 | COMPONENTS AND SUBASSEMBLIES |  |  |
| 14.1 | General |  |  |
|  | Where safety is involved, components and subassemblies meet relevant requirements | (see TABLE 1.A) |  |
| 14.2 | Motors | TBD-Info available |  |
| 14.2.1 | Motor temperatures |  |  |
|  | Does not present a hazard when stopped or prevented from starting; or | (see Forms A.1 and A.26B) |  |
|  | Protected by over-temperature or thermal protection device conform with 14.3 |  |  |
| 14.2.2 | Series excitation motors | Series excitation motors not present in equipment documentation. | N/A |
|  | Connected direct to device, if overspeeding causes a hazard | Critical overspeeding hazard information not available in documentation. | N/A |
| 14.3 | Overtemperature protection devices |  |  |
|  | Devices operating in a single fault condition | (see Form A.38) |  |
|  | - Reliable function is ensured | TBD-Info available |  |
|  | - rated to interrupt maximum current and voltage | TBD-Info available |  |
|  | - Does not operate in normal use | TBD-Info available |  |
|  | If self-resetting device used to prevent a hazard, protected part requires intervention before restarting | TBD-Info available |  |
| 14.4 | Fuse holders |  | P |
|  | No access to hazardous live parts | No hazardous live parts accessible per block diagram. | P |
| 14.5 | mains voltage selecting devices |  | P |
|  | Accidental change not possible | No evidence of accidental change possibility found. | P |
| 14.6 | mains transformers tested outside equipment | (see Forms A.39 and A.40) |  |
| 14.7 | Printed wiring boards |  |  |
|  | Data shows conformity with V-1 of IEC 60695-11-10 or better; or |  |  |
|  | Test shows conformity with V-1 of IEC 60695-11-10 or better | (see Form A.23) |  |
|  | Not applicable for printed wiring boards with limited-energy circuits (9.4) |  |  |
| 14.8 | Circuits used to limit transient overvoltages |  |  |
|  | Test conducted between each pair of mains supply terminals | (see Form A.41) |  |
|  | No ignition or overheating of other materials : | TBD-Info available | — |
|  | – no ignition | TBD-Info available |  |
|  | – no heat to other parts above the self-ignition points | TBD-Info not available |  |
|  | Safely suppressing and properly functional after applied tests | TBD-Info available |  |
| 15 | PROTECTION BY INTERLOCKS |  |  |
| 15.1 | General |  |  |
|  | Interlocks are designed to remove a hazard before operator exposed | TBD-Info available |  |
| 15.2 | Prevention of reactivation | TBD-Info not available |  |
| 15.3 | Reliability |  |  |
|  | Single fault unlikely to occur; or | TBD-Info not available |  |
|  | Cannot cause a hazard | TBD-Info not available |  |
| 16 | HAZARDS RESULTING FROM APPLICATION |  |  |
| 16.1 | reasonably foreseeable misuse |  | P |
|  | No hazards arising from settings not intended and not described in the instructions | Critical information on settings and instructions is missing. | P |
|  | Other cases of reasonably foreseeable misuse addressed by risk assessment | Risk assessment for misuse not documented in input files. | N/A |
| 16.2 | Ergonomic aspects |  |  |
|  | Factors giving rise to a hazard the risk assessment is reflecting those aspects: | Hazards include electric shock, overvoltage, and air pressure. | — |
|  | - limitation of body dimensions | Critical body dimension data not available in documentation. | N/A |
|  | - displays and indicators | Displays and indicators present: touchscreen, LEDs, monitor. | P |
|  | - accessibility and conventions of controls | Control types and locations not clearly documented. | P |
|  | - arrangement of terminals | Terminal arrangement clearly shown in block diagram. | P |
| 17 | RISK ASSESSMENT |  |  |
|  | risk assessment conducted, if hazard might arise and not covered by Clauses 6 to 16 |  |  |
|  | tolerable risk achieved by iterative documented process covering the following: |  | — |
|  | - risk analysis |  | N/A |
|  | Identifies hazards and estimates risk | Electric shock hazard mitigated by insulation and grounding. | P |
|  | - risk evaluation |  | P |
|  | Plan to judge acceptability of resulting risk level based on the estimated severity and likelihood of a risk | Risk level judged using insulation and clearance data. | P |
|  | - risk reduction |  | P |
|  | Initial risk reduced by counter measures; | Fuses, grounding, and insulation reduce electric shock risk. | P |
|  | Repeated risk evaluation without new risks introduced | No new risks introduced; repeated evaluation confirmed. | P |
|  | risks remaining after risk assessment addressed in instructions to responsible body: |  | — |
|  | Information contained how to mitigate these risks |  | P |
|  | Following principles in methods of risk reduction applied by manufacturer in given order: |  | — |
|  | - risks eliminated or reduced as far as possible | Risks reduced by fusing, grounding, insulation, and separation. | P |
|  | - Protective measures taken for risks that cannot be eliminated | Fuses and chassis grounding implemented for protection. | P |
|  | - User information about residual risk due to any defect of the protective measures | Residual risk information not documented in provided files. | P |
|  | Indication of particular training is required | No evidence of training indication found in documentation. | P |
|  | Specification of the need for personal protective equipment | No PPE requirement specified in documentation or images. | N/A |
|  | Conformity checked by evaluation of the risk assessment documentation | TBD-Info not available |  |
| ANNEX F | ROUTINE TESTS |  |  |
|  | Manufacturer ´s declaration |  |  |
| ANNEX H | QUALIFICATION OF CONFORMAL COATINGS FOR PROTECTION AGAINST POLLUTION |  |  |
| H.1 | General |  |  |
|  | Conformal coatings meet the requirements of Clause H.2 and H.3. |  |  |
| H.2 | Technical properties |  |  |
|  | Technical properties of conformal coatings are suitable for the intended application. In particular: |  | — |
|  | - Manufacturer indicate that it is a coating for PWBs; |  |  |
|  | - rated operating temperature include the temperature range of the indicated application; |  |  |
|  | - CTI, insulation resistance and dielectric strength are suitable for the intended application; |  |  |
|  | - Coating have adequate UV resistance, if it is exposed to sunlight; |  |  |
|  | - Flammability rating of the coating is at least the required flammability rating of the applied PWB. |  |  |
| H.3 | Qualification of coatings | (see Form A.42) |  |
|  | Coating complies with the conformity requirements. |  |  |
| ANNEX K | INSULATION REQUIREMENTS NOT COVERED BY CLAUSE 6.7 | (see Forms A.15 and A.18) |  |

| IEC 61010-1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  | Result — Remark |  | Verdict |
| 4.4 | TABLE: Testing in single fault condition – Results |  |  | Form A.1 |  |  |
| Test subclause | Fault No. | Fault description | Td 4.4.3 (note) |  | How was test terminated Comments | Meets 4.4.4 |
| note Td = Test duration in hh:mm:ss Record dielectric strength test on Form A.18 and temperature tests on Forms A.26A and / or A.26B. Record in the comments column for each test whether carried out during or after single fault condition. |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause |  | Requirement — Test |  |  |  |  |  | Result — Remark |  | Verdict |
| 5.1.3c) |  | TABLE: Mains supply |  |  |  |  |  | Form A.2 |  |  |
|  |  | Marked rating : |  |  |  |  |  | V |  | — |
|  |  | Phase : |  |  |  |  |  |  |  | — |
|  |  | Frequency : |  |  |  |  |  | Hz |  | — |
|  |  | Current : |  |  |  |  |  | A |  | — |
|  |  | Power : |  |  |  |  |  | W |  | — |
|  |  | Power : |  |  |  |  |  | VA |  | — |
| Test | Voltage |  | Frequency | Current |  | Power |  |  | Comments |  |
| No. | [V] |  | [Hz] | [A] |  | [W] | [VA] |  |  |  |
| NOTE – Measurements are only required for marked ratings. Initial inrush currents are not regarded. |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark |  | Verdict |
| 5.3 | TABLE: Durability of markings |  |  |  | Form A.3 |  |  |
| Marking method (see note) |  |  |  |  | Agent |  |  |
| 1) Adhesive label |  |  |  |  | A Water |  |  |
| 2) Ink printed |  |  |  |  | B Isopropyl alcohol 70% |  |  |
| 3) Laser marked |  |  |  |  | C (specify agent) |  |  |
| 4) Film-coated (plastic foil control panel) |  |  |  |  | D (specify agent) |  |  |
| 5) Imprinted on plastic (moulded in) |  |  |  |  | E (specify agent) |  |  |
| NOTE – Where applicable include print method, label material, ink or paint type, fixing method, adhesive and surface to which marking is fixed. |  |  |  |  |  |  |  |
| Marking location |  |  | Marking method (see above) |  |  |  |  |
| Identification (5.1.2) |  |  |  |  |  |  |  |
| Mains supply (5.1.3) |  |  |  |  |  |  |  |
| Fuses (5.1.4) |  |  |  |  |  |  |  |
| Terminals and operating devices (5.1.5.2) |  |  |  |  |  |  |  |
| Switches and circuit breakers (5.1.6) |  |  |  |  |  |  |  |
| Double/reinforced equipment (5.1.7) |  |  |  |  |  |  |  |
| Field wiring Terminal boxes (5.1.8) |  |  |  |  |  |  |  |
| Warning marking (5.2) |  |  |  |  |  |  |  |
| Battery charging (13.2.2) |  |  |  |  |  |  |  |
| Method | Test agent | Remains legible | Label loose | Curled edges |  | Comments |  |
|  |  | Verdict | Verdict | Verdict |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  | Result — Remark |  | Verdict |
| 6.2 | TABLE: List of accessible parts |  | Form A.4 |  |  |
| 6.1.2 | Exceptions |  |  |  | — |
| 6.2 | Determination of accessible parts |  |  |  | — |
| Item | Description | Determination method (note 5) |  | Exception under 6.1.2 (note 4) |  |
| NOTE 1 – Test fingers and pins are to be applied without force unless a force is specified (see 6.2.2) |  |  |  |  |  |
| NOTE 2 – Special consideration should be given to inadequate insulation and high voltage parts (see 6.2) |  |  |  |  |  |
| NOTE 3 – Parts are considered to be accessible if they could be touched in the absence of any covering which is not considered to provide suitable insulation (see 6.4). |  |  |  |  |  |
| NOTE 4 – Capacitance test may be required (see Form A.5). |  |  |  |  |  |
| NOTE 5 – The determination methods are: V = visual; R = rigid test finger; J = jointed test finger; P3 = pin 3 mm diameter; P4 = pin 4 mm diameter. |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  |  | Result — Remark |  |  |  |  |  | Verdict |
| 6 | TABLE: Values in normal condition |  |  |  |  |  |  | Form A.5 |  |  |  |  |  |  |
| 6.1.2 | Exceptions |  |  |  |  |  |  | 11.2 Cleaning and decontamination |  |  |  |  |  | — |
| 6.3.1 | Values in normal condition (see NOTE 1) |  |  |  |  |  |  | 11.3 Spillage |  |  |  |  |  | — |
| 6.6.2 | Terminals for external circuit |  |  |  |  |  |  | 11.4 Overflow |  |  |  |  |  | — |
| 6.10.3 | Plugs and connections |  |  |  |  |  |  |  |  |  |  |  |  | — |
| Item | Voltage |  |  | Current |  |  |  | Capacitance |  | 10 s / 5 s test (NOTE) |  |  | Comments |  |
| (see Form A.4) | V r.m.s. | V peak | V d.c. | Test circuit A1/A2/A3 | mA r.m.s. | mA peak | mA d.c. | C | mJ | V | C | mJ |  |  |
| NOTE – A 10 s test is specified in 6.1.2 a) b). A. 5 s test is specified in 6.10.3. The capacitance level versus voltage below the limits given from figure 3 of IEC 61010-1. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  |  | Result — Remark |  |  |  |  |  | Verdict |
| 6.3.2 | TABLE: Values in single fault condition |  |  |  |  |  |  | Form A.6 |  |  |  |  |  |  |
| Item | Subclause and | Voltage |  |  | Transient (see NOTE) |  | Current |  |  |  |  | Capacitance | Comments |  |
| (see Form A.4) | fault No. (see Form A.1) | V r.m.s. | V peak | V d.c. | V | s | Test circuit A1/A2/A3 |  | mA r.m.s. | mA peak | mA d.c. | F (see note) |  |  |
| NOTE – Transient voltages must be below the limits given from Figure 2 and the capacitance below the limits from figure 3 of IEC 61010-1. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  | Result — Remark |  | Verdict |
| 6.5.2.2 | TABLE: Cross-sectional area of bonding conductors Form A.7 |  |  |  |  |  |  |  |
| Conductor location |  | Cross-sectional area [mm²] |  |  |  |  |  | Verdict |
| Supplementary information: |  |  |  |  |  |  |  |  |
| 6.5.2.3 | TABLE: Tightening torque test Form A.8 |  |  |  |  |  |  |  |
| Conductor location |  |  |  |  | Size of screw |  | Tightening torque [Nm] | Verdict |
| Supplementary information: |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  | Result — Remark |  |  | Verdict |
| 6.5.2.4 | # TABLE: Bonding impedance of plug-connected equipment Form A.9 |  |  |  |  |  |  |  |  |  |
| accessible part under test |  |  | Test current [A] |  | Voltage attained after 1 min [V] |  |  | Calculated resistance (Maximum 0,1 or 0,2 ) [] (NOTE 1) |  | Verdict |
| NOTE 1 – For none-detachable power cord the impedance between protective conductor plug pin of mains cord and each accessible part shall not exceed 0,2 Ohm. |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |
| 6.5.2.5 | TABLE: Bonding impedance of permanently connected equipment Form A.10 |  |  |  |  |  |  |  |  |  |
| accessible part under test |  |  |  | Test current [A] |  | Voltage attained after 1 min (maximum 10 V) [V] |  |  |  | Verdict |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |
| 6.5.2.6 | # TABLE: Transformer protecive Bonding screen Form A.11 |  |  |  |  |  |  |  |  |  |
| accessible part under test |  | Test current (see note) [A] |  |  | Voltage attained after 1 min (maximum 10 V) [V] |  |  |  | Calculated resistance (maximum 0,1 ) [] | Verdict |
| NOTE – Test current must be twice the value of the overcurrent protection means of the winding. Test is specified in 6.5.2.6 a) or b). |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  | Result — Remark |  |  |  |  |  | Verdict |
| 6.5.4 | TABLE: Protective impedance |  |  |  |  | Form A.12 |  |  |  |  |  |  |
| A single component |  |  |  |  |  |  |  |  |  |  |  |  |
| Component |  | Location |  | Measured |  | Calculated | Rated |  |  | Verdict | Comments |  |
|  |  |  |  | Working voltage [V] | Current [A] | Power dissipation [W] | Working voltage [V] |  | Power dissipation [W] |  |  |  |
| A combination of components |  |  |  |  |  |  |  |  |  |  |  |  |
| Component |  |  | Location |  |  |  |  | Comments |  |  |  |  |
| NOTE – A protective impedance shall not be a single electronic device that employs electron conduction in a vacuum, gas or semiconductor. |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark |  |  |  |  | Verdict |
| 6.5.6 | TABLE: Current- or voltage-limiting device |  |  |  | Form A.13 |  |  |  |  |  |
| Component |  | Location | Measured |  |  | Rated |  | Verdict | Comments |  |
|  |  |  | Working voltage [V] | Current [A] |  | Working voltage [V] | Current [A] |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause |  |  | Requirement — Test |  |  |  |  |  |  |  | Result — Remark |  |  |  |  |  |  |  |  | Verdict |
| 6.7 |  |  | TABLE: Insulation requirements - Block diagram of system - Form A.14 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Pollution degree : |  |  |  |  |  |  |  |  |  |  | Overvoltage category : |  |  |  |  |  |  |  |  |  |
| Area | Location |  | Insulation type (note 1) |  |  | Working voltage |  |  | Clearance (note 3) |  | Creepage distance (note 3) |  |  |  |  |  | Test voltage | Comments (note 3) |  |  |
|  |  |  |  |  |  | RMS [V] | Peak [V] | Freq. [kHz] | [mm] |  | PWB [mm] | CTI |  | Other [mm] | CTI |  | (note 2) [V] |  |  |  |
| A |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| B |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| D |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| E |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| F |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| note 1 – Type of insulation: |  |  |  |  |  | note 2 - Types of voltage |  |  |  |  |  |  |  | note 3 - overvoltage categories |  |  |  |  |  |  |
| BI = Basic Insulation |  |  |  |  |  | Peak impulse test voltage (pulse) |  |  |  |  |  |  |  | or pollution degrees which differ |  |  |  |  |  |  |
| DI = Double Insulation |  |  |  |  |  | r.m.s. |  |  |  |  |  |  |  | should be shown under "Comments" |  |  |  |  |  |  |
| PI = Protective Impedance |  |  |  |  |  | d.c. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| RI = Reinforced Insulation |  |  |  |  |  | peak |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| SI = Supplementary Insulation |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| see also Form A.15 for further details |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary Information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause |  |  | Requirement — Test |  |  |  |  |  |  |  |  |  | Result — Remark |  |  |  |  |  |  |  | Verdict |
| 6.7 |  |  | TABLE: Insulation requirements - Clearances and Creepages |  |  |  |  |  |  |  |  |  | Form A.15 |  |  |  |  |  |  |  |  |
| 6.2.2 |  |  | Examination |  |  |  |  |  |  |  |  |  | 6.5.4 |  | Protective impedance |  |  |  |  |  | — |
| 6.4.2 |  |  | Enclosures and protective barriers |  |  |  |  |  |  |  |  |  | 6.5.6 |  | Current- or voltage-limiting device |  |  |  |  |  | — |
| 6.4.4 |  |  | Impedance |  |  |  |  |  |  |  |  |  | 9.6.1 |  | basic insulation between opposite polarity |  |  |  |  |  | — |
| Area | Location |  |  | Insulation type |  |  | Working Voltage (note 2) |  |  |  | Clearance |  |  |  | Creepage distance |  | CTI | Verdict | Comments |  |  |
|  | (See Form A.14) |  |  | (note 1) |  |  | RMS [V] |  | Peak [V] | Frequency [kHz] | Required [mm] |  | Measured [mm] |  | Required [mm] | Measured [mm] |  |  |  |  |  |
| A |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| B |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| D |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| E |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| F |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Note 1 – refer to Form A.14 for type of insulation shown in the insulation diagram Note 2 - to be used for definition of required insulation (see Form A.14) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Input supply voltage…….: |  |  |  |  | V |  |  | Hz |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause |  |  | Requirement — Test |  |  |  |  |  | Result — Remark |  |  |  |  |  |  |  | Verdict |
| 6.7 |  |  | TABLE: Insulation requirements - Clearances and Creepages |  |  |  |  |  | Form A.16 |  |  |  |  |  |  |  |  |
| 6.4.2 |  |  | enclosures or protective barriers |  |  |  |  |  | 9.6.1 |  | Overcurrent protection basic insulation between mains parts |  |  |  |  |  | — |
| 8 |  |  | Mechanical resistance to shock and impact |  |  |  |  |  | 10.5.1 |  | Integrity of clearances and creepage distances |  |  |  |  |  | — |
| Area | Location |  | Insulation type | Mechanical tests (NOTE) |  |  |  |  |  |  | Test at max. | Measured after test (if required) |  | Verdict | Comments |  |
|  | (See Form A.14) |  |  | Applied force | Rigidity (8.2) |  | Drop (8.3) |  |  |  | rated ambient | Clearance | Creepage distance |  |  |  |
|  |  |  |  | [N] | Static (8.2.1) | Impact (8.2.2) | Normal (8.3.1) |  | Hand-held/ Plug-in |  | (10.5.1) | [mm] | [mm] |  |  |  |
| A |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| B |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| D |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| E |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| F |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE – Refer to Form A.18 for dielectric strength tests following the above tests. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement – Test |  |  |  |  |  |  | Result — Remark |  |  |  |  | Verdict |  |  |  |
| 6.7.2.2.2 | TABLE: Reliability of potted components |  |  |  |  |  |  | Form A.17 (optional) |  |  |  |  |  |  |  |  |
| 14.1 b) | Components and subassemblies |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Temperature Cycling Test |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Manufacturer : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Type : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Construction : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Potting compound : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Creepage distances measured : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Clearances measured : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Thickness through insulation : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Adhesive test Pass/Fail : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test temperature T °C : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Cycles at U= AC 500 V |  |  |  |  |  |  |  |  | Leakage current (at AC 500 V) mA |  |  |  |  |  |  |  |
| Number of cycles |  |  |  | Date |  |  |  |  | 68 h / |  | 1 h / | 2 h / |  | 1 h / |  |  |
|  |  |  |  |  |  |  |  |  | 125 °C |  | 25 °C | 0 °C |  | 25 °C |  |  |
| 1. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 2. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 3. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 4. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 5. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 6. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 7. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 8. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 9. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| 10. Cycle from |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |
| After Cycling Test : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Humidity conditioning |  |  |  |  |  |  |  |  | 48 h |  |  |  |  |  |  |  |
| Requirements for dielectric strength (s. insulation diagram) |  |  |  |  |  |  |  |  | Test voltage V r.m.s. |  |  |  |  | Verdict |  |  |
| Basic insulation ______ V r.m.s. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary insulation ______ V r.m.s. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Reinforced insulation ______ V r.m.s. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE - to be used for evaluation of components containing insulation through solid insulation, when the component standard require thermal cycling test. Ref Clause 14.1 and Figure 15, option b) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  | Result — Remark | Verdict |
| 6.8 | TABLE: Dielectric strength tests |  |  |  |  | Form A.18 |  |
| 4.4.4.1 b) | Conformity after application of single fault conditions1 |  |  |  |  |  |  |
| 6.4 | Primary means of protection2 |  |  |  |  |  |  |
| 6.6 | Connections to external circuits |  |  |  |  |  |  |
| 6.7 | Insulation requirements2 (see Annex K) |  |  |  |  |  |  |
| 6.10.2 | Fitting of non-detachable mains supply cords1 |  |  |  |  |  |  |
| 9.2 a) 2) | Eliminating or reducing the sources of ignition within the equipment |  |  |  |  |  |  |
| 9.4 c) | Limited-energy circuit |  |  |  |  |  |  |
| 9.6.1 | Overcurrent protection basic insulation between mains - parts |  |  |  |  |  |  |
|  | Test site altitude : |  |  |  |  | m | — |
|  | Test voltage correction factor (see table 10) : |  |  |  |  |  | — |
| Location or references from Forms A.1 and A.14 |  | Clause or sub-clause | Humidity | Working voltage | Test voltage | Comments (note) | Verdict |
|  |  |  | Yes/No | [r.m.s./d.c.] | [r.m.s./peak/d.c.] |  |  |
| 1 Record the fault, test or treatment applied before the dielectric strength test. 2 Humidity preconditioning required. |  |  |  |  |  |  |  |
| NOTE: Test duration may be recorded. |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark |  |  |  |  |  | Verdict |
| 6.10.2 | TABLE: Cord anchorage |  |  |  | Form A.19 |  |  |  |  |  |  |
| Location |  | Mass [kg] | Pull [N] | Verdict |  | Torque [Nm] | Verdict |  | Comment |  |  |
| Dielectric strength test for 1 min. (6.8.3.1) : |  |  |  |  |  |  |  | V r.m.s. |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  | Result — Remark |  |  |  |  |  |  |  |  |  |  | Verdict |
| 7. | TABLE: Protection against mechanical HAZARDS |  |  |  |  |  | Form A.20 |  |  |  |  |  |  |  |  |  |  |  |
| 7.3.4 | Limitation of force and pressure |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |
| 7.3.5 | Gap limitations between moving parts |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |
| Part / Location |  | Clause 7.3.4 |  | Clause 7.3.5.1 |  |  |  |  |  |  |  |  | Clause 7.3.5.2 |  |  | Verdict | Comments |  |
|  |  | Continuous | Temporary | Minimum gaps [mm] |  |  |  |  |  |  |  |  | Maximum gaps [mm] |  |  |  |  |  |
|  |  | Contact pressure max. 50 N /cm² @ max. 150 N | max. 250 N / 3 cm² @ max. 0,75 s | Torso 500 | Head 300 | Leg 180 |  | Foot 120 | Toes 50 | Arm 120 | Hand 100 | Finger 25 | Head 120 | Foot 35 | Finger 4 |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement – Test |  |  |  |  |  |  |  | Result - Remark |  |  |  |  |  | Verdict |
| 7.4 | TABLE: Stability |  |  |  |  |  |  |  | Form A.20A |  |  |  |  |  |  |
|  | Equipment height / mass : |  |  |  |  |  |  |  |  |  | mm |  |  | kg | — |
|  | Equipment (Containers) loaded : |  |  |  |  |  |  |  | [yes / no] |  |  |  |  |  | — |
|  | Castors at unfavourable position : |  |  |  |  |  |  |  | [yes / no] |  |  |  |  |  | — |
|  | Doors, drawers and movable arms closed : |  |  |  |  |  |  |  | [yes / no] |  |  |  |  |  | — |
|  | Doors and drawers at unfavourable position : |  |  |  |  |  |  |  | [yes / no] |  |  |  |  |  | — |
| Location |  |  | Tilt angle | Applied force |  |  |  |  |  | Comments |  |  |  |  | Verdict |
|  |  |  | 10° | 250 N | 20% [N] |  | 800 N | 4 times load [N] |  |  |  |  |  |  |  |
| Front side |  |  |  |  |  |  | — |  |  |  |  |  |  |  |  |
| Left side |  |  |  |  |  |  | — |  |  |  |  |  |  |  |  |
| Rear side |  |  |  |  |  |  | — |  |  |  |  |  |  |  |  |
| Right side |  |  |  |  |  |  | — |  |  |  |  |  |  |  |  |
| Top side |  |  | — |  |  |  |  |  |  |  |  |  |  |  |  |
| Working surface |  |  | — | — | — |  |  |  |  |  |  |  |  |  |  |
| Ledge |  |  | — | — | — |  |  |  |  |  |  |  |  |  |  |
| Castor / support foot |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Castor / support foot removed |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 7.6 | TABLE: Wall mounting |  |  |  |  |  |  |  | Form A.20B |  |  |  |  |  |  |
|  | Equipment weight : |  |  |  |  |  |  |  |  |  |  | kg |  |  | — |
|  | Equipment mounted as specified by manufacturer : |  |  |  |  |  |  |  | [yes / no] |  |  |  |  |  | — |
|  | Equipment mounted at plasterboard (drywall) : |  |  |  |  |  |  |  | [yes / no] |  |  |  |  |  | — |
|  | More than one fastener used : |  |  |  |  |  |  |  | [yes / no] |  |  |  |  |  | — |
|  | Test maintained (after 5 s to 10 s to full load) : |  |  |  |  |  |  |  | 1 min |  |  |  |  |  | — |
| Location |  | Applied weight |  |  |  |  |  |  | Comments |  |  |  |  |  | Verdict |
|  |  | 4 times weight [kg] |  |  |  | 2 times weight [kg] |  |  |  |  |  |  |  |  |  |
| Mounting brackets |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement – Test | Result - Remark |  |  |  | Verdict |
| 8.2 | TABLE: Enclosure rigidity test | Form A.21A |  |  |  |  |
| 8.2.1 | Static test |  |  |  |  |  |
|  | Material of enclosure : | Metal / non-metallic |  |  |  | — |
|  | Preparation for the test: |  |  |  |  | — |
|  | Operated at ambient temperature : |  | °C |  | h | — |
| Location |  | Comments |  |  |  | Verdict |
| 1) |  |  |  |  |  |  |
| 2) |  |  |  |  |  |  |
| 3) |  |  |  |  |  |  |
| 4) |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |
| 8.2.2 | TABLE: Impact test |  |  |  |  |  |
|  | Material of enclosure : | Metal / non-metallic |  |  |  | — |
|  | Corresponding IK-code : |  |  |  |  | — |
|  | Preparation for the test: |  |  |  |  | — |
|  | Cooled to (temperature) : |  |  | °C |  | — |
| Location |  | Comments |  |  |  | Verdict |
| 1) Top |  |  |  |  |  |  |
| 2) Side left / right |  |  |  |  |  |  |
| 3) Bottom |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement – Test |  |  | Result - Remark |  | Verdict |
| 8.3 | TABLE: Drop test |  |  | Form A.21B |  |  |
| 8.3.1 | Other equipment |  |  |  |  |  |
| Location |  | Raised up to |  | Comments |  | — |
|  |  | [mm] | 30 ° |  |  |  |
| 1) |  |  |  |  |  |  |
| 2) |  |  |  |  |  |  |
| 3) |  |  |  |  |  |  |
| 4) |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |
| 8.3.2 | Hand-held equipment and direct plug-in equipment |  |  |  |  |  |
|  | Material of enclosure : |  |  | Metal / non-metallic |  | — |
|  | Preparation for the test: |  |  |  |  | — |
|  | Cooled to (temperature) : |  |  |  | °C | — |
| Location |  |  |  | Comments |  | Verdict |
| 1) Side |  |  |  |  |  |  |
| 2) Edge |  |  |  |  |  |  |
| 3) Corner |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Clause |  | Requirement — Test |  | Result — Remark | Verdict |
| 9 | TABLE: Protection against the spread of fire |  |  | Form A.22 |  |
| Item | Source of hazard or area of the equipment considered (circuit, component, liquid etc.) |  | Protection Method (9.1 a, b or c) | Protection details | Verdict |
| Supplementary information: |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark |  |  |  | Verdict |  |
| 9.3.2 | TABLE: Constructional requirements |  |  |  | Form A.23 |  |  |  |  |  |
| 14.7 | Printed wiring boards |  |  |  |  |  |  |  |  |  |
| Material tested : |  |  |  |  |  |  |  |  | — |  |
| Generic name : |  |  |  |  |  |  |  |  | — |  |
| Material manufacturer : |  |  |  |  |  |  |  |  | — |  |
| Type : |  |  |  |  |  |  |  |  | — |  |
| Colour : |  |  |  |  |  |  |  |  | — |  |
| Conditioning details : |  |  |  |  |  |  |  |  | — |  |
|  |  |  | Sample |  |  |  |  |  |  |  |
|  |  |  | 1 | 2 |  | 3 | 4 | 5 |  | 6 |
| Thickness of specimen |  | mm |  |  |  |  |  |  |  |  |
| Duration of flaming after first Application |  | s |  |  |  |  |  |  |  |  |
| Duration of flaming plus glowing After second application |  | s |  |  |  |  |  |  |  |  |
| Specimen burns to holding clamp |  | Yes/No |  |  |  |  |  |  |  |  |
| Cotton ignited |  | Yes/No |  |  |  |  |  |  |  |  |
| Sample result |  | Pass/Fail |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark |  |  |  | Verdict |
| 9.4 | TABLE: Limited-energy circuit |  |  |  | Form A.24 |  |  |  |  |
| Item |  | 9.4 a) | 9.4 b) Current limitation (NOTE) |  |  | 9.4 c) | Decision | Comments |  |
| or Location (see Form A.22) |  | Maximum potential in circuit voltage r.m.s./d.c. [V] | Maximum available current [A] | Overload protection after 120 s [A] |  | Circuit separation | Yes/No |  |  |
| NOTE – Maximum values see Tables 17 and 18 of IEC 61010-1 |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  | Result — Remark |  | Verdict |
| 9.5 | TABLE: Requirements for equipment containing or using flammable liquids Form A.25 |  |  |  |  |
| Type of liquid |  | 9.5 Flammable liquids |  |  | Verdict |
|  |  | b) Quantity |  | c) Containment |  |
| Supplementary information: |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  |  |  | Result — Remark |  |  |  |  |  | Verdict |
| 10. | TABLE: Temperature Measurements |  |  |  |  |  |  |  | Form A.26A |  |  |  |  |  |  |
| 10.1 | Surface temperature limits – normal condition and / or single fault condition |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 10.2 | Temperature of windings – normal condition and / or single fault condition |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 10.3 | Other temperature measurements |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Operating conditions: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Frequency : |  |  | Hz |  | Test room ambient temperature (ta) : |  |  |  |  |  |  |  | °C |  |  |
| Voltage : |  |  | V |  | Test duration : |  |  |  |  |  |  | h |  | min |  |
| Part / Location |  |  |  | tm [C] |  | tc [C] |  | tmax [C] |  | Verdict | Comments |  |  |  |  |
| NOTE 1 - tm = measured temperature |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| tc = tm corrected (tm–ta+ 40 C or max. rated ambient) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| tmax = maximum permitted temperature |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 2 - see also 14.1 with reference to component operating conditions |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 3 - Record values for normal condition and / or single fault condition in this Form use additional form if necessary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 4 - see Form A.26B for details of winding temperature measurements |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  |  |  | Result — Remark |  |  |  |  |  |  | Verdict |
| 10.2 | TABLE: Temperature of windings Resistance method Temperature Measurements |  |  |  |  |  |  |  | Form A.26B |  |  |  |  |  |  |  |
| 4.4.2.7 | Mains transformers |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 14.2.1 | Motor temperatures |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Operating conditions : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Frequency : |  |  | Hz | Test room ambient temperature (ta1/ta2) : |  |  |  |  |  |  |  | / |  | °C (initial / final) |  |  |
| Voltage : |  |  | V | Test duration : |  |  |  |  |  |  |  |  | h |  | min |  |
| Part / Designation |  | Rcold [] |  | Rwarm [] | Current [A] | tr [K] |  | tc [C] |  | tmax [C] | Verdict |  | Comments |  |  |  |
| NOTE 1- Rcold = initial resistance |  |  |  |  |  |  | Rwarm = final resistance |  |  |  |  |  |  |  |  |  |
| tr = temperature rise |  |  |  |  |  |  | tc = tr corrected (tc= tr + [40 C or max rated ambient]) |  |  |  |  |  |  |  |  |  |
| tmax = maximum permitted temperature |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 2 - Indicate insulation class (IEC 60085) under comments (optional) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 3 - Record values for normal condition and / or single fault condition in this Form use additional form if necessary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  | Result — Remark |  |  |  | Verdict |
| 10.5.2 | TABLE: Resistance to heat of non-metallic enclosures |  |  | Form A.27 |  |  |  |
|  | Test method used: |  |  |  |  |  | — |
|  | Non-operative treatment : |  | [ ] |  |  |  |  |
|  | Empty enclosure : |  | [ ] |  |  |  |  |
|  | Operative treatment : |  | [ ] |  |  |  |  |
|  | Temperature during tests : |  |  |  |  |  | — |
| Description |  | Material | Comments |  |  |  | Verdict |
| Dielectric strength test (6.8) : |  |  |  |  | V | [r.m.s./peak/d.c.] |  |
| NOTE – Within 10 minutes of the end of treatment suitable tests in acc. to 8.2 and 8.3 must be conducted and pass criteria of 8.1. |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark |  | Verdict |
| 10.5.3 | TABLE: Insulating material |  |  |  | Form A.28 |  |  |
| 10.5.3 1) | Ball-pressure test |  |  |  |  |  |  |
|  | Max. allowed impression diameter : |  |  |  | 2 mm |  | — |
| Part |  | Test temperature [°C] |  | Impression diameter [mm] |  |  | Verdict |
| Supplementary information: |  |  |  |  |  |  |  |
| 10.5.3 2) | Vicat softening test (ISO 306) |  |  |  |  | Form A.29 |  |
| Part |  |  | Vicat softening temperature [°C] |  |  | Thickness of sample [mm] | Verdict |
| Supplementary information: |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  |  | Result — Remark |  |  |  |  |  |  | Verdict |
| 8 | TABLE: Mechanical resistance to shock and impact |  |  |  |  |  |  | Form A.30 |  |  |  |  |  |  |  |
| 11 | Protection against hazards from fluids and solid foreign objects |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Voltage tests can be carried out once after performing the tests of clause 8 and clause 11. However, if voltage tests are carried out separately after each set of tests, two forms can be used. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Clause 8 tests |  |  |  | Clause 11 tests |  |  |  |  |  |  |  |  |  |
| Location (see Form A.14) |  | Static (8.2.1) 30 N | Impact (8.2.2) | Normal (8.3.1) | Handheld Plug-in (8.3.2) | Cleaning (11.2) | Spillage (11.3) |  | Overflow (11.4) | IEC 60529 (11.6) | Working voltage [r.m.s./d.c.] | Test voltage [r.m.s./peak/d.c.] | Verdict | Comments |  |
| NOTE – Use r.m.s., d.c. or peak to indicate the used test voltage. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause |  | Requirement — Test |  |  |  |  |  |  |  |  |  | Result — Remark |  |  |  |  |  | Verdict |  |
| 11.7.2 |  | TABLE: Leakage and rupture at high pressure |  |  |  |  |  |  |  |  |  |  |  | Form A.31 |  |  |  |  |  |
| Part |  |  | Maximum permissible working pressure [MPa] |  |  | Test pressure [MPa] |  |  |  | Leakage Yes / No |  |  | Deformation Yes / No |  | Burst Yes / No | Comments |  |  |  |
| NOTE – see also Annex G with requirements for USA and Canada. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 11.7.3 |  | TABLE: Leakage from low-pressure parts |  |  |  |  |  |  |  |  |  |  |  | Form A.32 |  |  |  |  |  |
| Part |  |  |  |  | Test pressure [MPa] |  |  | Leakage Yes / No |  |  |  | Comments |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  | Result — Remark | Verdict |
| 12.2.1 | TABLE: Ionizing radiation |  |  | Form A.33 |  |
| 12.2.1.2 | Equipment intended to emit radiation |  |  |  |  |
| Locations tested |  | Measured values [µSv/h] | Verdict | Comments |  |
| Supplementary information: |  |  |  |  |  |
| 12.2.1.3 | Equipment not intended to emit radiation |  |  | Form A.34 |  |
|  | Max. allowed effective dose rate at 100 mm………: |  |  | 1 µSv/h | — |
| Locations tested |  | Measured values [µSv/h] | Verdict | Comments |  |
| Supplementary information: |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark | Verdict |
| 12.5.1 | TABLE: Sound level |  |  |  | Form A.35 |  |
| Locations tested |  | Measured maximum sound pressure level dB(A) |  | Calculated maximum sound power level |  |  |
| At operator’s normal position and at bystanders’ positions |  |  |  |  |  |  |
| a) |  |  |  |  |  |  |
| b) |  |  |  |  |  |  |
| c) |  |  |  |  |  |  |
| d) |  |  |  |  |  |  |
| e) |  |  |  |  |  |  |
| f) |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |
| 12.5.2 | TABLE: Ultrasonic pressure |  |  |  | Form A.36 |  |
| Locations tested |  | Measured values |  | Comments |  |  |
|  |  | [dB] | [kHz] |  |  |  |
| At operator’s normal position |  |  |  |  |  |  |
| At 1 m from the enclosure |  |  |  |  |  |  |
| a) |  |  |  |  |  |  |
| b) |  |  |  |  |  |  |
| c) |  |  |  |  |  |  |
| d) |  |  |  |  |  |  |
| e) |  |  |  |  |  |  |
| NOTE – No limit is specified at present, but a limit of 110 dB above the reference pressure value of 20 Pa is under consideration for applicable frequencies between 20 kHz and 100 kHz. |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  | Result — Remark |  | Verdict |
| 13.2.2 | TABLE: Batteries and battery charging |  | Form A.37 |  |  |
|  | Battery load and charging circuit diagram: |  |  |  |  |
|  | Battery type : |  |  |  | — |
|  | Battery manufacturer/model/catalogue No. : |  |  |  | — |
|  | Battery ratings : |  |  |  | — |
|  | Reverse polarity instalment test |  |  |  |  |
| Single component failures |  | Verdict |  |  |  |
| Component |  | Open circuit |  | Short circuit |  |
| Supplementary information: |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  | Result — Remark | Verdict |
| 14.3 | TABLE: Overtemperature protection devices |  |  |  | Form A.38 |  |
| Reliability test |  |  |  |  |  |  |
| Component |  | Type (note) | Verdict | Comments |  |  |
| NOTE: |  |  |  |  |  |  |
| NSR = non-self-resetting (10 times) |  |  |  |  |  |  |
| NR = non-resetting (1 time) |  |  |  |  |  |  |
| SR = self-resetting (200 times) |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause |  | Requirement — Test |  |  |  |  |  |  |  |  | Result — Remark |  |  |  | Verdict |  |
| 4.4.2.7 |  | TABLE: Mains transformer |  |  |  |  |  |  |  |  | Form A.39 |  |  |  |  |  |
| 4.4.2.7.2 |  | Short circuit |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 14.6 |  | Mains transformers tested outside equipment |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Type : |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |
| Manufacturer : |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |
| Test in equipment |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test on bench |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test repeated inside equipment (see 14.6) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Optional – Insulation class (IEC 60085) of the lowest rated winding : |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |
| Winding identification |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Type of Protector for winding (NOTE 1) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Elapsed time |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Current, A primary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| secondary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Winding temperature, °C primary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| (see NOTE 2) secondary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Tissue paper / cheesecloth OK ? (Pass / Fail) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Voltage tests (see NOTE 3) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Primary to secondary |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |
| Primary to core |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |
| Secondary to secondary |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |
| Secondary to core |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |
| Verdict |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 1: | Primary fuse |  |  |  |  |  |  | - PF / | ( ) |  |  | A |  |  |  |
|  | Secondary fuse |  |  |  |  |  |  | - SF / | ( ) |  |  | A |  |  |  |
|  | Overtemperature protection |  |  |  |  |  |  | - OP / | ( ) |  |  | °C |  |  |  |
|  | Impedance protection |  |  |  |  |  |  | - Z |  |  |  |  |  |  |  |
| NOTE 2: | Indicate method of measurement |  |  |  |  |  |  | - TC = with thermocouple |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | - R = resistance method |  |  |  |  |  |  |  |
|  | If resistance method is used, record resistance in cold and warm condition in Form A.26B. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 3: | Record the voltage applied and the type of voltage (r.m.s. / d.c. / peak) and for |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | results use NB = no breakdown or B = breakdown |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause |  | Requirement — Test |  |  |  |  |  |  |  |  | Result — Remark |  |  |  |  | Verdict |  |
| 4.4.2.7 |  | TABLE: Mains transformer |  |  |  |  |  |  |  |  | Form A.40 |  |  |  |  |  |  |
| 4.4.2.7.3 |  | Overload tests (for mains transformers) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 14.6 |  | Mains transformers tested outside equipment |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Type : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |
| Manufacturer : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |
| Test in equipment |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test on bench |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test repeated inside equipment (see 14.6) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Optional – Insulation class (IEC 60085) of the lowest rated winding : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |
| Winding identification |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Type of Protector for winding (NOTE 1) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Elapsed time |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Current, A primary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| secondary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Winding temperature, °C primary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| (see NOTE 2) secondary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Tissue paper / cheesecloth OK ? (Pass / Fail) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Voltage tests (see NOTE 3) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Primary to secondary |  |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |  |
| Primary to core |  |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |  |
| Secondary to secondary |  |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |  |
| Secondary to core |  |  |  | ______ |  | V | ____ |  |  |  |  |  |  |  |  |  |  |
| Verdict |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 1: | Primary fuse |  |  |  |  |  |  | - PF / | ( ) |  |  | A |  |  |  |  |  |
|  | Secondary fuse |  |  |  |  |  |  | - SF / | ( ) |  |  | A |  |  |  |  |  |
|  | Overtemperature protection |  |  |  |  |  |  | - OP / | ( ) |  |  | °C |  |  |  |  |  |
|  | Impedance protection |  |  |  |  |  |  | - Z |  |  |  |  |  |  |  |  |  |
| NOTE 2: | Indicate method of measurement |  |  |  |  |  |  | TC = with thermocouple |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | R = resistance method |  |  |  |  |  |  |  |  |  |
|  | If resistance method is used, record resistance in cold and warm condition in Form A.26B. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE 3: | Record the voltage applied and the type of voltage (r.m.s. / d.c. / peak) and for |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | results use NB = no breakdown or B = breakdown |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  |  |  |  | Result — Remark |  |  |  |  |  |  | Verdict |
| 14.8 | TABLE: Circuits used to limit transient overvoltages |  |  |  |  |  |  |  | Form A.41 |  |  |  |  |  |  |  |
| Circuit / Designation |  | Overvoltage Category | Mains voltage [V r.m.s.] |  |  | Test voltage [V] | tm [C] | tc [C] |  | tmax [C] | Ignited Yes / No | Safely suppressed Yes / No | Properly functional Yes / No | Verdict | Comments |  |
| Test room ambient temperature: |  |  |  |  | °C |  |  |  |  |  |  |  |  |  |  |  |
| NOTE - tm = measured temperature |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| tc = tm corrected (tm–ta+ 40 C or max. rated ambient) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| tmax = maximum permitted temperature |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Conformity is checked by applying 5 positive and 5 negative impulses with the applicable impulse withstand voltage, spaced up to 1 min apart, from a hybrid impulse generator (see IEC 61180-1). |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement – Test |  |  |  |  |  |  |  |  | Result — Remark |  |  |  |  | Verdict |  |  |  |
| Annex H | TABLE: Qualification of conformal coating for protection against pollution |  |  |  |  |  |  |  |  | Form A.42 |  |  |  |  |  |  |  |  |
| Technical properties |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Manufacturer : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |  |  |
| Type : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | — |  |  |  |
| Meet requirements of ANSI / UL 746E : |  |  |  |  |  | [yes / no] |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Manufacturer declaration of coating material : |  |  |  |  |  | [yes / no] |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Operating temperature of coating : |  |  |  |  |  | [ ] °C |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comparative tracking index (CTI) : |  |  |  |  |  | [ ] |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Insulation resistance : |  |  |  |  |  | [ ] MΩ |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Dielectric strength : |  |  |  |  |  | [ ] V |  |  |  |  |  |  |  |  |  |  |  |  |  |
| UV resistance (if required) : |  |  |  |  |  | [yes / no] |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Flammability rating : |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Preparation of the test specimens conducted : |  |  |  |  |  | [yes / no] |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Item |  | Test conditioning |  |  | Parameter | Td |  | Samples |  |  |  |  |  |  | Verdict | Comments |  |  |  |
|  |  |  |  |  |  | h |  | 1 | 2 | 3 | 4 | 5 |  | 6 |  |  |  |  |  |
| 1 |  | Cold |  |  |  | 24 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  | Dry heat |  |  |  | 48 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  | Rapid temp. change |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  | Damp heat |  |  |  | 24 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 |  | Adhesion of coating |  |  | 5 N |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Visual inspection |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 6 |  | Humidity |  |  |  | 48 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 7 |  | Insulation resistance |  |  | ≥ 100 MΩ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Visual inspection |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NOTE Td = Test duration time |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Supplementary information: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement – Test |  |  | Result — Remark |  | Verdict |
|  | TABLE: Additional or special tests conducted |  |  |  | Form A.43 |  |
| Clause and name of test |  | Test type and condition | Observed results |  |  | — |
| Supplementary information: |  |  |  |  |  |  |

| IEC 61010-1 |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  |  |  |  | Result — Remark |  |  |  |  | Verdict |
|  | table 1.A: List of components and circuits relied on for safety |  |  |  |  |  |  |  |  |  |  |
| Unique component reference or location |  | Application/function | Manufacturer / trademark (note 1) |  | Type / model |  | Technical data (note 2) |  | Standard | Mark(s) of conformity evidence of acceptance (NOTE 3 and 4) |  |
| NOTE 1 List all different manufacturers of the above components |  |  |  | 4 asterisk indicates mark assuring agreed level of surveillance |  |  |  |  |  |  |  |
| 2 May include electrical, mechanical values |  |  |  |  |  |  |  |  |  |  |  |
| 3 List licence no or method of acceptance |  |  |  |  |  |  |  |  |  |  |  |

List of test equipment used:

A completed list of used test equipment shall be provided in the Test Reports when a Customer’s Testing Facility according to CTF stage 1 or CTF stage 2 procedure has been used.

Note: This page may be removed when CTF stage 1 or CTF stage 2 are not used. See also clause 4.8 in 
OD 2020 for more details.

| Clause | Measurement / testing | Testing / measuring equipment / material used, (Equipment ID) | Range used | Last Calibration date | Calibration due date |
| --- | --- | --- | --- | --- | --- |

Statement of Measurement Uncertainty

The Test Report shall include a statement concerning the uncertainty of the measurement systems used for the tests conducted when it is required by the standard, client or other authorities.

In such cases, the table below is to be used for reporting U of M.

This page may be removed from the final Test Report when not required. See also clause 4.8 in 
OD 2020 for more details.

| Clause # | Parameter/ Measurement / test method | Requirement % or k | Calculated U of M* |
| --- | --- | --- | --- |

*Note: Calculations leading to the reported value are on file with the NCB