## Sheet1
Component Name | Category | Confidence | Source Types | Evidence Count | Image URLs | Image Captions | Guide Reference | critical | confidence | rules_passed | rules_passed_count | rule_score | reasoning
Infinity MTx device | Assembly | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Assembly name alone does not indicate direct safety function, hazardous circuit involvement, or enclosure role; insufficient evidence for criticality per IEC 61010.
User Interface Touchscreen | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Touchscreen is a user interface; no evidence it is part of safety or hazardous circuitry, nor does it maintain spacings or prevent access.
Sample Loading Tray | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Sample loading tray is a mechanical component; no indication it encloses hazardous circuitry or is required by IEC 61010 for safety.
Barcode Scanner | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Barcode scanner is an identification device; not located in safety or hazardous circuitry, nor required for electrical safety compliance.
Exhaust Air Vents | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Exhaust air vents may relate to thermal management, but no evidence they enclose hazardous circuitry or are required by IEC 61010 for electrical safety.
Enclosure | Assembly | high | guide | 4 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 4] | 2 | 0.25 | The enclosure is required by IEC 61010 to prevent access to hazardous voltages and parts, and is specifically relied upon to enclose hazardous circuitry. Its removal or modification would directly compromise user safety.
Power Supply | Component | high | guide | 3 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 3] | 2 | 0.25 | The power supply is required by the standard and is located in hazardous circuitry, as it interfaces with mains voltage and provides power to the system. Its compliance and integrity are essential for electrical safety.
Main PCB | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.5 | [] | 0 | 0 | The main PCB is referenced only in the guide with low-medium confidence and no evidence of direct safety or hazardous circuit involvement. Without further information, it cannot be classified as critical.
MPA (Microfluidic Processing Assembly) | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.5 | [] | 0 | 0 | The MPA subsystem is referenced with low-medium confidence and no evidence of safety or hazardous circuit involvement. No rules are satisfied based on available data.
Optimizer Kit | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.5 | [] | 0 | 0 | The Optimizer Kit is referenced with low-medium confidence and no evidence of safety or hazardous circuit involvement. No rules are satisfied based on available data.
Infinity MTx Platform | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | No evidence provided that the Infinity MTx Platform subsystem is directly required by IEC 61010, located in safety/hazardous circuitry, or fulfills any criticality rules. Subsystem designation alone is insufficient for criticality.
Infinity Microfluidic Processing Assembly (MPA) | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | No evidence that the Microfluidic Processing Assembly is required by IEC 61010 or involved in safety/hazardous circuitry. No indication of enclosure, spacing, or abnormal operation relevance.
Sterilized Membranes | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Sterilized membranes are not shown to be required by IEC 61010, nor do they appear to be involved in electrical safety, hazardous circuitry, or critical segregation functions.
Actuating Lid | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Actuating lid is not demonstrated to enclose hazardous circuitry or be required by IEC 61010. No evidence of safety or spacing function provided.
Microfluidic Channels | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Microfluidic channels are not shown to be involved in electrical safety, hazardous circuitry, or required by IEC 61010. No criticality rules satisfied based on available data.
Chevron Ridge | Component | high | guide | 2 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | No evidence that the Chevron Ridge is required by IEC 61010, located in safety/hazardous circuitry, or used for spacings/segregation. It appears to be a mechanical guide feature with no direct safety function.
MPA Lid Latch | Mechanical Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | The MPA Lid Latch is a mechanical component; no evidence it encloses hazardous circuitry or is required by IEC 61010. Unless it prevents access to hazardous parts, it is not critical by the rules.
MPA Lid | Enclosure Component | high | guide | 2 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [4] | 1 | 0.125 | The MPA Lid is an enclosure component likely preventing access to hazardous circuitry, satisfying rule 4. This is a typical requirement under IEC 61010 for operator protection.
Air Inlet Port | Fluidic Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | The Air Inlet Port is a fluidic component with no evidence of direct safety or hazardous circuit involvement. Not required by IEC 61010 or used for segregation.
Air Outlet Port | Fluidic Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | The Air Outlet Port is a fluidic component; no evidence it is involved in safety or hazardous circuitry, nor required by IEC 61010. Not critical by the rules.
Filter Membrane (0.2 um) | Filtration Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Filter membrane is a fluidic filtration component, not required by IEC 61010, not located in safety or hazardous electrical circuitry, and does not maintain electrical spacings or prevent access to hazardous voltages.
Collection Chamber | Assembly | high | guide | 2 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | Collection chamber is an assembly for fluid collection, not required by IEC 61010, not part of safety or hazardous electrical circuitry, and does not maintain electrical spacings or prevent access to hazardous voltages.
Collection Chamber Lid | Enclosure Component | high | guide | 3 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | Collection chamber lid is an enclosure component for fluid containment, not required by IEC 61010, not part of safety or hazardous electrical circuitry, and does not maintain electrical spacings or prevent access to hazardous voltages.
Collection Chamber Latch | Mechanical Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Collection chamber latch is a mechanical component for securing the chamber, not required by IEC 61010, not part of safety or hazardous electrical circuitry, and does not maintain electrical spacings or prevent access to hazardous voltages.
Sample Chamber | Fluidic Component | high | guide | 2 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | Sample chamber is a fluidic component, not required by IEC 61010, not part of safety or hazardous electrical circuitry, and does not maintain electrical spacings or prevent access to hazardous voltages.
Microfluidic Chip | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Microfluidic chip is a subsystem for fluid handling; no evidence it is required by IEC 61010, located in safety/hazardous circuitry, or used for electrical segregation. No indication of special evaluation or hazard if changed.
Microfluidic Channel | Fluidic Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Microfluidic channel is a fluidic path, not referenced in IEC 61010 as critical for electrical safety. No evidence of safety/hazardous circuit location, segregation, or hazard prevention.
Compressed Air Inlet | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Compressed air inlet is a pneumatic interface, not directly covered by IEC 61010 electrical safety requirements. No evidence of safety/hazardous circuit involvement or hazard prevention.
Push-to-Connect Fitting | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Push-to-connect fitting is a mechanical fluidic connector, not required by IEC 61010 for electrical safety. No evidence of safety/hazardous circuit location or segregation function.
Tubing | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Tubing is a fluidic transport component, not referenced in IEC 61010 for electrical safety. No evidence of safety/hazardous circuit involvement or hazard prevention.
Air Intake Fan | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Air intake fan is not directly required by IEC 61010, nor does it enclose, segregate, or form part of safety/hazardous circuitry. Unless specifically used for safety-related thermal management, it is not critical.
Power Connection Port | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [1, 3, 4] | 3 | 0.375 | Power connection port is required by IEC 61010 for safe connection to mains, is located in hazardous (mains) circuitry, and prevents direct access to hazardous voltages.
Power Cord (C13 to 5-15) | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [1, 3, 4, 7] | 4 | 0.5 | Power cord is required by IEC 61010, is part of hazardous (mains) circuitry, prevents access to live parts, and is relied upon for safe abnormal operation (e.g., cord integrity).
Power Switch | Component | high | guide | 2 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 2, 3, 7] | 4 | 0.5 | Power switch is required by IEC 61010, is located in safety and hazardous circuitry, and is relied upon for safe abnormal operation (disconnection from mains).
USB-A Connection Port | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | USB-A connection port is not required by IEC 61010, nor does it enclose, segregate, or form part of safety/hazardous circuitry unless used for safety functions, which is not indicated.
LAN Connection Port | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | LAN Connection Port is not typically required by IEC 61010, nor does it directly interface with hazardous or safety circuitry. No evidence of criticality per rules.
Hinge | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Hinge is not specified as enclosing or preventing access to hazardous circuitry, nor is it required by the standard. No evidence of criticality per rules.
Sample Chambers | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Sample Chambers are not indicated as being part of hazardous or safety circuitry, nor do they maintain spacings or segregation. No evidence of criticality per rules.
Lid Latch | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Lid Latch is not shown to enclose hazardous circuitry or be required by the standard. No evidence of criticality per rules.
0.2 um Membrane Filters | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | 0.2 um Membrane Filters are not required by IEC 61010 for electrical safety and do not interact with hazardous or safety circuitry. No evidence of criticality per rules.
Internal Electrical Components | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.6 | [3] | 1 | 0.125 | Internal electrical components are likely located in hazardous circuitry per IEC 61010 definitions. Any change or failure may expose users to electrical hazards.
Channel | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.4 | [] | 0 | 0 | The term 'Channel' is ambiguous and lacks evidence of safety or hazardous function. No rules can be confirmed without further technical detail.
Collection Chamber Latches | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.6 | [4] | 1 | 0.125 | Collection chamber latches may prevent access to hazardous areas or substances, satisfying enclosure/access prevention requirements.
Ventilation Slots | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.6 | [4] | 1 | 0.125 | Ventilation slots can provide access to internal hazardous circuitry and must be evaluated for finger/access protection per IEC 61010.
Rear Connections | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.6 | [3] | 1 | 0.125 | Rear connections typically interface with hazardous voltages or circuits, making them critical per hazardous circuitry requirements.
Connection Lines and Cables | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.7 | [3, 5] | 2 | 0.25 | Connection lines and cables may be located in hazardous circuitry and are used to maintain spacings or segregation, both of which are critical per IEC 61010.
Compressed Air Connection | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.5 | [] | 0 | 0 | Compressed air connection is not directly referenced by IEC 61010 for electrical safety and does not enclose, segregate, or interact with hazardous electrical circuitry.
Compressor | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.5 | [] | 0 | 0 | Compressor subsystem is not directly involved in electrical safety per IEC 61010 unless it interfaces with hazardous electrical circuits, which is not indicated by the provided data.
Electrical Outlet | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [1, 3] | 2 | 0.25 | Electrical outlets are required by IEC 61010 and are located in hazardous circuitry, making them critical for electrical safety compliance.
Attached Air Lines | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.5 | [] | 0 | 0 | Attached air lines are not referenced by IEC 61010 for electrical safety and do not interact with hazardous electrical circuits or provide segregation.
Infinity Instrument | Subsystem | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | No evidence provided that the Infinity Instrument subsystem is directly required by IEC 61010, located in safety/hazardous circuitry, or relied upon for safety. Insufficient information to classify as critical.
Compressed Air Cylinder | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.7 | [8] | 1 | 0.125 | Compressed Air Cylinder may present a physical hazard if changed or deleted, such as pressure release or mechanical failure. IEC 61010 requires evaluation of non-electrical hazards; thus, it is critical.
Infinity Carrying Case | Assembly | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Infinity Carrying Case is an external assembly with no evidence of enclosing hazardous circuitry or maintaining safety-related spacings. Not critical per IEC 61010.
Foam Layers | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Foam Layers are not shown to be required for safety, segregation, or enclosure of hazardous circuitry. No evidence of criticality per IEC 61010.
Instrument Screen | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.3 | [] | 0 | 0 | Instrument Screen is not indicated as part of safety or hazardous circuitry, nor required for maintaining spacings or enclosure. Not critical per IEC 61010.
Instrument Casing | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [4] | 1 | 0.125 | Instrument casing typically encloses hazardous circuitry and prevents user access to live parts, satisfying rule 4. This is a fundamental requirement for IEC 61010 compliance.
Tray | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | Tray is not generally required by IEC 61010, nor does it typically enclose hazardous circuitry or maintain spacings. No evidence it is part of safety or hazardous circuits.
Rubber Feet | Component | low-medium | guide | 1 |  |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Rubber feet do not enclose hazardous circuitry, are not required by IEC 61010, and are not relied upon for safety or abnormal operation. No critical safety function identified.