## Risk Assessment-Operator
Risk # | Sub-System | Risk | Description | Severity | Probablility | Pre-Mitigation PRN | Mitigation | Post-Mitigation Probabliity | Post-Mitigation PRN | Verfication Reference | Risk Acceptable ? Y/N
op-1 | Drawer | Pinch | Operator leaves fingers in drawer as it closes | Minor | Rare | 1 | Drawer designed to not cause pain or pinching to finger | Rare | 1 | add 61010 report #
op-2 | Drawer | Bio Hazard | Contamination from operator error inputting biological substance into instrument | Moderate | Unlikely | 2 | Consumable has filters installed. Drawer is designed to contain leaked biological substance to prevent contamination of rest of instrument. | Unlikely | 2 | add 61010 report #
op-3 | System | Shock | Faulty wiring/electrical connections leading to part of the chassis being electrified | Severe | Unlikely | 3 | Grounding scheme complies with 61010 safety | Unlikely | 3 | add 61010 report #
op-4 | System | Shock | Faulty designed structure leading to exposed wiring or possibility of electrical shock. | Severe | Unlikely | 3 | All high voltage connections protected and in complience with 61010 safety | Unlikely | 3 | add 61010 report #
op-5 | System | Tripping | Power entry and bulkhead locations that are not positioned correctly | Moderate | Rare | 2 | Plug and Power Entry locations designed in easy to see/reach areas and not near top of instrument | Unlikely | 1
op-6 | Pneumatic System | Release of pressure | Release of high pressure causing destruction or damage. Projection of damaged debris under pressure | Moderate | Unlikely | 2 | All parts were spec’d with double the pressure rating and reviewed during safety testing | Unlikely | 2 | add 61010 report #
op-7 | System | Fire caused by short in system | Electrical short causes fire | Catastrophic | Unlikely | 2 | primary side fuses limit current drawn from the wall | Unlikely | 2 | add 61010 report # & Lab equipment CE standard 61326
op-8 | System | Damage to external equipment | Device emissions cause damage to other equipment | Minor | Rare | 1 | Product Designed and tested to 61010 EMC / EMI testing standard | Rare | 1 | add 61010 report #
op-9 | System | Crush | Operator pulls unit over | Severe | Rare | 2 | Feet are placed wide on the instrument, most of the unit weight is located on the base. | Rare | 2 | add 61010 report #
op-10 | System | fire | blocked ventilation or restricted ventilation | Catastrophic | Unlikely | 2 | User manual has install information, System evaluated under blocked airflow conditions | Unlikely | 2 | add 61010 report #
op-11 | Pneumatic System | fire | wrong gas (oxygen) attached to system | Catastrophic | Rare | 2 | air line has gas specfic connector (air) | Unlikely | 2 | add 61010 report #
Note on H3: Tim Etchells: can the user replace the filter ? if so is there risk there too?

## Risk Assessment-Service
Risk # | Sub-System | Risk | Description | Severity | Probablility | Pre-Mitigation PRN | Mitigation | Post-Mitigation Probabliity | Post-Mitigation PRN | Verfication Reference | Risk Acceptable ? Y/N
Svc-1 | System | Shock | Shock from touching live part of system with covers removed | Catastrophic | Likely | 3 | Service manual warnings, service personnel are trained, and connections are covered. | Rare | 2 | training documentation/ 61010 report
Svc-2 | System | Pinch hazard while performing maintenance | While servicing the internal system, unit activates | Severe | Possible | 2 | Service manual warnings, service personnel are trained, interlocks must be over-ridden (personnel are trained to use EMO stop for panel interlock override) and use temporary covers while moving gantry. Firmware has a position error timeout that will remove power to gantry in the event of a blockage.(under force limit for 61010). | Rare | 2 | training documentation
Svc-3 | System | Crush | Service operator pulls unit over while installing | Severe | Unlikely | 2 | instructions to tie down system and secure in field service manual | Unlikely | 2 | user manual/
Svc-4 | Drawer | Bio Hazard | Contamination from operator error inputting biological substance into instrument | Moderate | Unlikely | 2 | Consumable has filters installed. Drawer is designed to contain leaked biological substance to prevent contamination of rest of instrument. | Unlikely | 2 | test report ?
Srvc-5 | Pneumatic System | Release of pressure | Release of high pressure causing destruction or damage during service | Moderate | Unlikely | 2 | Service manual warnings, service personnel are trained to de presurise before service | Unlikely | 2 | add 61010 report #
Svc-6 | System | Bio Hazard | System not decontaminated prior to service | Moderate | Unlikely | 2 | Decontamination protocol in place prior to service or return | Unlikely | 2 | user manual/

## Risk_tables
Do not move this table, its used as the reference into the table for the RPN fomula
Probablility
[5 empty] | Frequent | Likely | Possible | Rare | Unlikely
|  |  | Severity | Catastrophic | 3 | 3 | 3 | 2 | 2
|  |  |  | Severe | 3 | 3 | 2 | 2 | 1
|  |  |  | Moderate | 3 | 2 | 1 | 1 | 1
|  |  |  | Minor | 2 | 1 | 1 | 1 | 1
KEY | Catagory | Description |  | Severity of harm |  |  |  |  | Likelyhood
1 | Broadly acceptable | This fulfils the requirement for TOLERABLE RISK. |  | Catastrophic | One or more fatalities | System or facility loss | Chemical release with acute or public health impact |  | Frequent | More than 5 times a year
2 | As low as reasonably practicable | This does not automatically fulfil the requirement for TOLERABLE RISK. If possible, these RISKS should be reduced further to Category 1. If not possible, then the instructions should contain a description of the RISK so that the RESPONSIBLE BODY can take appropriate steps to protect the safety of OPERATORS. |  | Severe | Disabling injury/illness | Major subsystem loss or facility damage | Chemical release with temporary environmental or public health impact |  | Likely | More than once per year, but not more than five times a year
3 | Intolerable | This contains RISKS that are not TOLERABLE RISKS. |  | Moderate | Medical treatment or restricted work activity | Minor subsystem loss or facility damage | Chemical release triggering external reporting requirements |  | Possible | More than once in five years, but not more than one a year
|  |  |  | Minor | First aid only | Non-serious equipment or facility damage | Chemical release requiring only routine cleanup without reporting |  | Rare | More than once in ten years, but not more than one in five years
[9 empty] | Unlikely | No more than once in ten years