## Sheet1
Component Name | Category | Confidence | Source Types | Evidence Count | Image URLs | Guide Reference | critical | confidence | rules_passed | rules_passed_count | rule_score | reasoning
AC Inlet | Connector | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 4] | 3 | 0.375 | The AC inlet is required by IEC 61010 for mains connection, is located in hazardous (mains) circuitry, and encloses/prevents access to hazardous voltages. Its integrity is essential for user safety.
Fuse | Protection Device | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 2, 3, 7] | 4 | 0.5 | The fuse is required by IEC 61010 for overcurrent protection, is part of safety circuitry, located in hazardous mains circuit, and is relied upon for safe abnormal operation (overcurrent/short circuit protection).
Chassis Ground | Grounding | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 2, 7, 8] | 4 | 0.5 | Chassis ground is required by IEC 61010 for protective earthing, is part of safety circuitry, is relied upon for safe abnormal operation (faults), and its removal would create a shock hazard.
SWPS (RPS-500-24-C) | Power Supply | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 7] | 3 | 0.375 | The power supply is required by the standard, is located in hazardous circuitry (mains input), and is relied upon for safe abnormal operation (e.g., insulation, fault protection).
PRI (Primary Side) | Power Supply Section | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | The primary side of the power supply is hazardous (mains voltage) and is used to maintain spacings/segregation from secondary circuits, as required by IEC 61010.
SEC (Secondary Side) | Power Supply Section | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3] | 1 | 0.125 | The secondary side of a power supply is considered hazardous circuitry under IEC 61010 if it can present hazardous energy or voltage. It must be evaluated for insulation, spacings, and protection.
DC 24V Output | Power Output | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 7] | 2 | 0.25 | The DC 24V output is part of hazardous circuitry if it can present hazardous energy, and is relied upon for safe abnormal operation (e.g., overvoltage, short-circuit protection).
Secondary Circuits (PCB) | Electronic Circuit | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | Secondary circuits on the PCB may be hazardous and are used to maintain spacings/segregation from primary circuits, as required by IEC 61010.
Metal Enclosure | Enclosure | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 4] | 2 | 0.25 | The metal enclosure is required by IEC 61010 to prevent access to hazardous circuitry and provide protection against electric shock.
INLET MANIFOLD | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 1 | [] | 0 | 0 | The inlet manifold is a pneumatic component and does not interact with electrical safety or hazardous electrical circuitry per IEC 61010.
PRESSURE SENSOR SPT25-10-0200A | Sensor | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [2, 3, 7] | 3 | 0.375 | Pressure sensor is likely used in safety-related and hazardous pneumatic circuitry to monitor and control pressure, and may be relied upon for safe abnormal operation (e.g., overpressure detection).
NC VALVE 181168 | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [2, 3, 7] | 3 | 0.375 | NC (normally closed) valve is typically located in hazardous and safety-related pneumatic circuits, controlling access to pressurized gas and relied upon for safe abnormal operation (e.g., fail-safe shutoff).
MECHANICAL PREREGULATOR | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.85 | [3, 7] | 2 | 0.25 | Mechanical preregulator is located in hazardous pneumatic circuitry and is relied upon for safe abnormal operation by limiting upstream pressure.
MAIN MANIFOLD | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.85 | [3, 8] | 2 | 0.25 | Main manifold is part of hazardous pneumatic circuitry; removal or modification could result in uncontrolled gas flow, presenting a hazard.
E-REGULATOR ITV0050-3BL | Pneumatic Regulator | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [2, 3, 7] | 3 | 0.375 | E-regulator controls pressure in hazardous pneumatic circuits, is part of safety control, and is relied upon for safe abnormal operation (e.g., pressure regulation failure could cause hazard).
PRESSURE RELIEF SAFETY VALVE (125 PSI) | Safety Device | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 2, 3, 7, 8] | 5 | 0.625 | Pressure relief safety valves are required by IEC 61010 for overpressure protection in pneumatic/hydraulic systems. They are located in safety circuitry, directly prevent hazardous overpressure, and are relied upon for safe abnormal operation. Removal or failure would result in a significant hazard.
PARTICULATE FILTER | Filter | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.8 | [] | 0 | 0 | Particulate filters are not specifically required by IEC 61010 and do not directly prevent access to hazardous energy or maintain electrical safety spacings. Unless used in safety or hazardous circuitry, they are generally non-critical for electrical safety compliance.
NO VENT VALVE 126419 | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [3, 4, 8] | 3 | 0.375 | A vent valve in a pneumatic system can enclose or prevent access to hazardous pressure and is located in hazardous circuitry. Its removal or malfunction could result in uncontrolled release of pressure, presenting a safety hazard.
AIR BALAST AVT-32-16 | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | An air ballast is typically used for pressure stabilization and is not directly required by IEC 61010 for safety, nor does it prevent access to hazardous energy or maintain electrical spacings. It is not critical unless specifically used in safety or hazardous circuitry.
MAC VALVE (VENT VALVE) | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [3, 4, 8] | 3 | 0.375 | A vent valve in a pneumatic system is located in hazardous circuitry and can prevent access to hazardous pressure. Its removal or failure could result in a hazardous condition, making it critical for safety compliance.
PISTON MANIFOLD | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | The piston manifold is a pneumatic component and, based on available evidence, does not directly interface with hazardous electrical circuitry or provide electrical safety functions required by IEC 61010. No rules are satisfied without further evidence of safety function or hazard containment.
THERMISTOR 10K 235-1483-ND | Sensor | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | The thermistor is a sensor, typically used for temperature measurement. Unless it is part of a safety circuit or relied upon for abnormal operation protection, it is not critical per IEC 61010. No evidence provided of such use.
PNEUMATIC VALVE DOWN 181168 | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | The pneumatic valve (DOWN) is a control element in the pneumatic system. Without evidence of its role in safety or hazardous circuit containment, it does not meet any criticality rules.
PNEUMATIC VALVE UP 181168 | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | The pneumatic valve (UP) is similar to the DOWN valve; it is not shown to be part of safety or hazardous electrical circuitry, nor does it maintain spacings or containment. No rules are satisfied.
PNEUMATIC COMPRESSION SYSTEM | Pneumatic System | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | The pneumatic compression system is a subsystem for pneumatic operation. No evidence is provided that it encloses, prevents access to, or interfaces with hazardous electrical circuitry per IEC 61010. No rules are satisfied.
2X CYLINDER SWITCH M9P | Switch | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | No evidence that the 2X CYLINDER SWITCH M9P is required by IEC 61010, located in safety/hazardous circuitry, or relied upon for safety. No indication it maintains spacings or prevents access to hazardous parts. No special evaluation or hazard identified if changed or deleted.
DISTRIBUTION MANIFOLD | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | The DISTRIBUTION MANIFOLD is a pneumatic component and not directly referenced by IEC 61010 for electrical safety. No evidence it is part of safety/hazardous circuitry, maintains spacings, or prevents access to hazardous electrical parts. No special evaluation or hazard identified.
VALVE 6724 / 299249 | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | VALVE 6724 / 299249 is a pneumatic valve, not directly covered by IEC 61010 electrical safety requirements. No evidence it is required by standard, part of safety/hazardous circuitry, or maintains spacings. No special evaluation or hazard identified.
8 PORT CARTRIDGE | Cartridge Holder | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | 8 PORT CARTRIDGE is a cartridge holder, not referenced by IEC 61010 for electrical safety. No evidence it is required by standard, part of safety/hazardous circuitry, or maintains spacings. No special evaluation or hazard identified.
PRESENCE SENSOR | Sensor | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | PRESENCE SENSOR is a sensor, but no evidence it is required by IEC 61010, located in safety/hazardous circuitry, or relied upon for safety. No indication it maintains spacings or prevents access to hazardous parts. No special evaluation or hazard identified.
CARTRIDGE HOLDER AND DRAWER SUB ASSEMBLY | Mechanical Assembly | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [4] | 1 | 0.125 | The cartridge holder and drawer sub assembly likely encloses or prevents access to hazardous circuitry or substances, satisfying rule 4. This is a common requirement for mechanical assemblies in IEC 61010 to prevent user contact with hazards.
DRAWER MOTOR | Motor | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | The drawer motor, as described, does not appear to be located in safety or hazardous circuitry, nor is it relied upon for safety functions or abnormal operation. No evidence of criticality per IEC 61010 rules.
BASEPLATE AND SKINS | Mechanical Structure | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [4, 5] | 2 | 0.25 | Baseplate and skins typically form the enclosure, preventing access to hazardous parts and maintaining required spacings/segregation, satisfying rules 4 and 5. These are critical for compliance with IEC 61010.
TOUCHSCREEN | User Interface | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.8 | [] | 0 | 0 | Touchscreen is a user interface component and, unless it provides isolation or safety functions, is not critical per IEC 61010. No evidence it is required for safety or hazard prevention.
BARCODE SCANNER | Peripheral Device | high | guide, image | 2 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.8 | [] | 0 | 0 | Barcode scanner is a peripheral device and does not appear to be involved in safety or hazardous circuitry, nor is it required by the standard for safety. No rules satisfied.
NFC READER | Peripheral Device | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.8 | [] | 0 | 0 | NFC Reader is a peripheral device, not typically involved in safety or hazardous circuitry, nor required by IEC 61010 unless used for safety functions. No evidence of criticality.
USB HUB | Electronic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.8 | [] | 0 | 0 | USB HUB is an electronic component for data/power distribution, not directly involved in safety or hazardous circuits per IEC 61010 unless used for safety isolation. No evidence of criticality.
MAIN PCB | Printed Circuit Board | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | MAIN PCB is likely to contain hazardous circuitry and is relied upon for maintaining spacings/segregation per IEC 61010. Critical for electrical safety.
DISCRETE ELEC COMPONENT | Electronic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | DISCRETE ELEC COMPONENT is generic; without evidence of safety or hazardous function, cannot be deemed critical under IEC 61010.
SUB ASSEMBLY | Assembly | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | SUB ASSEMBLY is generic; unless it encloses hazardous circuitry or is part of safety function, not critical per IEC 61010. No evidence provided.
PCB | Printed Circuit Board | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | PCB is likely to carry hazardous voltages and is relied upon to maintain required spacings and segregation per IEC 61010. Any failure or modification could compromise safety barriers.
INTEL NUC CH-42-1.2 | Computer | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.8 | [] | 0 | 0 | Computer (INTEL NUC) is not directly involved in safety or hazardous circuitry based on available evidence. No indication it is relied upon for safety or abnormal operation.
3X USB PASSTHRU 1698-01C | Connector | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.8 | [] | 0 | 0 | USB passthru connector is not located in hazardous or safety circuitry and does not maintain spacings or segregation. No evidence of criticality per IEC 61010.
POWER POWER ENTRY MODULE FPYSSA106FBR | Power Entry Module | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 4] | 3 | 0.375 | Power entry module is required by IEC 61010, located in hazardous mains circuitry, and encloses/prevents access to hazardous voltages. Its integrity is essential for user safety.
POWER SUPPLY RPS-500-SF | Power Supply | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 7] | 3 | 0.375 | Power supply is required by standard, located in hazardous circuitry, and is relied upon for safe abnormal operation (e.g., overcurrent, overvoltage protection).
FAN 1688-1554-ND | Cooling Fan | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [3, 7, 8] | 3 | 0.375 | Cooling fans are critical if they prevent overheating of hazardous circuitry, which may lead to fire or electric shock. Their removal or failure can result in unsafe abnormal operation and other hazards.
FAN 1053-1243-ND | Cooling Fan | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [3, 7, 8] | 3 | 0.375 | As with other cooling fans, this component is relied upon to maintain safe temperatures in hazardous circuits. Its failure or removal may result in unsafe abnormal operation and other hazards.
LOW FLOW SENSOR SFC5500-50SCCM | Flow Sensor | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | Flow sensors are generally not directly involved in safety or hazardous circuitry unless specifically used for safety interlocks or hazardous gas detection. No evidence of such use is provided.
HIGH FLOW SENSOR SFC5500-0.5SLM | Flow Sensor | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | FALSE | 0.7 | [] | 0 | 0 | High flow sensors are not inherently critical unless used in safety interlocks or hazardous detection. No evidence of such application is provided.
BACKPLANE | Mechanical/Electrical Interface | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.95 | [5, 8] | 2 | 0.25 | Backplane is relied upon for maintaining spacings and segregation of circuits, which is essential for electrical safety. Its removal or modification may introduce hazards.
Infinity MTx Platform | Assembly | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | The Infinity MTx Platform is an assembly; no evidence provided that it directly encloses hazardous circuitry, is part of safety or hazardous circuits, or is required by IEC 61010. No specific safety function identified.
Infinity Microfluidic Processing Assembly (MPA) | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | The Microfluidic Processing Assembly is a subsystem; no evidence it is part of safety or hazardous circuitry, nor required by IEC 61010. No indication of enclosure or segregation function.
Sterilized Membranes | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.8 | [] | 0 | 0 | Sterilized membranes are not electrical safety components and do not interact with hazardous voltages or circuits per IEC 61010. No evidence of safety or segregation function.
Actuating Lid | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Actuating lid may provide mechanical function, but no evidence it encloses hazardous circuitry or is required for electrical safety per IEC 61010. No safety or segregation function identified.
Microfluidic Channels | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.8 | [] | 0 | 0 | Microfluidic channels are not part of electrical safety or hazardous circuitry, nor do they maintain spacings or segregation of electrical circuits per IEC 61010. No safety function identified.
Chevron Ridge | Component | high | guide | 2 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | No evidence that the Chevron Ridge is required by IEC 61010, located in safety/hazardous circuitry, or used for spacings/segregation. No indication of special evaluation or abnormal operation reliance.
Enclosure | Component | high | guide | 3 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 4] | 2 | 0.25 | Enclosure is required by IEC 61010 to prevent access to hazardous circuitry and is specifically referenced in the standard for protection against electric shock and mechanical hazards.
Infinity MTx device | Assembly | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | The Infinity MTx device as an assembly is not a single component relied upon for safety; criticality must be assessed at the sub-component level.
Front Panel | Enclosure | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [4] | 1 | 0.125 | Front Panel is part of the enclosure, preventing access to hazardous circuitry, which is a requirement under IEC 61010.
User Interface Touchscreen | Human-Machine Interface | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | User Interface Touchscreen is not directly required by IEC 61010 for safety, nor does it enclose or segregate hazardous circuitry. No evidence of special evaluation or abnormal operation reliance.
Sample Loading Tray | Mechanical Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Sample Loading Tray is a mechanical guide component and does not enclose, segregate, or interact with hazardous or safety circuitry per available evidence.
Exhaust Air Vents | Ventilation | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Exhaust Air Vents provide ventilation but, with current evidence, do not enclose hazardous circuitry or maintain required spacings; no direct safety function identified.
Main Body Enclosure | Enclosure | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [4] | 1 | 0.125 | Main Body Enclosure prevents access to hazardous circuitry, directly satisfying enclosure requirements per IEC 61010.
MPA (Microfluidic Processing Assembly) | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | MPA (Microfluidic Processing Assembly) is a subsystem with no evidence of direct safety or hazardous circuit involvement; insufficient data for criticality.
Optimizer Kit | Accessory/Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | Optimizer Kit is an accessory/subsystem with no evidence of safety or hazardous circuit function; not required by IEC 61010 for safety.
Power Supply | Electrical Component | high | guide | 4 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 3, 7] | 3 | 0.375 | Power supply is required by IEC 61010, located in hazardous circuitry, and relied upon for safe abnormal operation (e.g., overvoltage, overcurrent protection).
Printed Circuit Board (PCB) | Electrical Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [3, 5] | 2 | 0.25 | PCB is located in hazardous circuitry and used to maintain spacings/segregation per IEC 61010 requirements.
Internal Instrumentation | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | Internal instrumentation is a subsystem; no direct evidence it is required by standard, in safety/hazardous circuitry, or relied upon for safety.
MPA Lid Latch | Mechanical Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | MPA Lid Latch is a mechanical component; no evidence it encloses hazardous circuitry or is required for safety per IEC 61010.
MPA Lid | Enclosure Component | high | guide | 2 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [4] | 1 | 0.125 | MPA Lid is an enclosure component that prevents access to hazardous circuitry, satisfying IEC 61010 enclosure requirements.
Air Inlet Port | Fluidic Interface | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Air inlet port is a fluidic interface and does not directly interact with hazardous electrical circuitry or safety functions per IEC 61010. No evidence of safety or hazardous circuit involvement.
Air Outlet Port | Fluidic Interface | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Air outlet port is a fluidic interface, not required by IEC 61010 for electrical safety, and not involved in hazardous or safety circuitry. No evidence of criticality.
Filter Membrane (0.2 um) | Filtration Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Filter membrane (0.2 um) is a filtration component for fluidic separation, not required by IEC 61010 for electrical safety, and not involved in hazardous or safety circuitry.
Collection Chamber | Assembly | high | guide | 3 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.8 | [] | 0 | 0 | Collection chamber is an assembly for fluid collection, not enclosing hazardous electrical circuitry or required for electrical safety per IEC 61010.
Collection Chamber Lid | Enclosure Component | high | guide | 2 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.8 | [] | 0 | 0 | Collection chamber lid is an enclosure component for fluid containment, not enclosing or preventing access to hazardous electrical circuitry per IEC 61010.
Collection Chamber Latch | Mechanical Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [4] | 1 | 0.125 | The Collection Chamber Latch prevents access to internal hazardous circuitry or fluids, satisfying rule 4. Its removal or failure could expose users to hazards.
Sample Chamber | Fluidic Subsystem | high | guide | 2 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [3, 4] | 2 | 0.25 | The Sample Chamber contains hazardous fluids and may enclose hazardous circuitry, satisfying rules 3 and 4. It is critical for user protection and containment.
Microfluidic Chip | Microfluidic Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | The Microfluidic Chip is not directly referenced as enclosing, preventing access, or being part of hazardous or safety circuitry per available evidence.
Microfluidic Channel | Microfluidic Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.6 | [] | 0 | 0 | The Microfluidic Channel is not shown to be part of hazardous or safety circuitry, nor does it enclose or segregate hazardous areas per available documentation.
Compressed Air Inlet | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [3] | 1 | 0.125 | The Compressed Air Inlet interfaces with a hazardous energy source (compressed air), satisfying rule 3. Failure or modification could introduce safety hazards.
Push-to-Connect Fitting | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Push-to-connect fitting is typically used for pneumatic connections and does not directly interface with electrical safety or hazardous circuitry per IEC 61010. No evidence of safety or hazardous function.
Tubing | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Tubing is generally used for fluid or air transport and is not required by IEC 61010 for electrical safety unless it encloses hazardous circuits, which is not indicated here.
Air Intake Fan | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Air intake fan is not directly involved in safety or hazardous electrical circuitry per IEC 61010 unless used for forced ventilation of hazardous areas, which is not specified.
Power Connection Port | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.95 | [1, 3, 4] | 3 | 0.375 | Power connection port is required by IEC 61010, interfaces with hazardous voltage, and prevents direct access to hazardous circuitry. It is critical for electrical safety.
Power Cord (C13 to 5-15) | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.95 | [1, 3, 4] | 3 | 0.375 | Power cord is required by IEC 61010, interfaces with hazardous voltage, and prevents direct access to hazardous circuitry. It is critical for electrical safety.
Power Switch | Component | high | guide | 2 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 3, 4, 7] | 4 | 0.5 | Power switch is required by IEC 61010, is located in hazardous (mains) circuitry, encloses/prevents access to hazardous voltages, and is relied upon for safe abnormal operation (emergency shutoff).
USB-A Connection Port | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.8 | [] | 0 | 0 | USB-A port is typically low voltage and not part of safety or hazardous circuitry per IEC 61010; no evidence of criticality or special evaluation.
LAN Connection Port | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.8 | [] | 0 | 0 | LAN port is generally low voltage, not in safety or hazardous circuits, and not required by IEC 61010 for safety; no evidence of criticality.
Lid Latch | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [4] | 1 | 0.125 | Lid latch prevents access to hazardous circuitry, satisfying enclosure/access prevention per IEC 61010.
Membrane Filter (0.2 um) | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Membrane filter is not required by IEC 61010 for electrical safety, not in hazardous or safety circuitry, and does not maintain spacings or segregation.
Sample Chambers | Assembly | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [4] | 1 | 0.125 | Sample chambers may enclose or prevent access to hazardous circuitry or substances, satisfying rule 4. IEC 61010 requires barriers to hazardous areas.
Internal Electrical Components | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [3] | 1 | 0.125 | Internal electrical components are located in hazardous circuitry per IEC 61010 definitions, satisfying rule 3. These require compliance with safety requirements.
Channel | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.5 | [] | 0 | 0 | No evidence that the 'Channel' is required by the standard, located in safety/hazardous circuitry, or used for segregation. Insufficient data for criticality.
Collection Chamber Latches | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.7 | [4] | 1 | 0.125 | Collection chamber latches may prevent access to hazardous areas, satisfying rule 4. IEC 61010 requires secure barriers to hazardous parts.
Ventilation Slots | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [4, 5] | 2 | 0.25 | Ventilation slots can prevent access to hazardous circuitry and maintain required spacings/segregation, satisfying rules 4 and 5 per IEC 61010.
Rear Connections | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [3, 4] | 2 | 0.25 | Rear connections may provide access to hazardous voltages or circuits and often serve as entry points to hazardous areas, thus enclosure and prevention of access is required by IEC 61010.
Connection Lines and Cables | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [3, 5] | 2 | 0.25 | Connection lines and cables can carry hazardous voltages/currents and are relied upon to maintain required spacings and segregation between circuits per IEC 61010.
Compressed Air Connection | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.7 | [8] | 1 | 0.125 | Compressed air connection, if changed or deleted, may introduce mechanical or pneumatic hazards, such as uncontrolled release of pressure, which is addressed by IEC 61010 for non-electrical hazards.
Compressor | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.7 | [8] | 1 | 0.125 | Compressor is a source of mechanical and pneumatic hazard; removal or modification could result in unsafe operation, covered under IEC 61010 for non-electrical hazards.
Compressed Air Line | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.7 | [8] | 1 | 0.125 | Compressed air line, if altered, may result in hazardous pneumatic conditions (e.g., burst, uncontrolled release), which IEC 61010 requires to be controlled.
Infinity Instrument | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [1, 3] | 2 | 0.25 | As a subsystem, the Infinity Instrument is required by the standard for evaluation and may contain hazardous circuitry; thus, it is critical.
Carrying Case | Assembly | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | The carrying case is not required by IEC 61010, does not enclose hazardous circuitry, nor is it relied upon for safety; it is non-critical.
Foam Padding | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Foam padding is not required by the standard, does not maintain spacings, nor does it prevent access to hazardous circuitry; it is non-critical.
Instrument Screen | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Instrument screen is not required for safety, does not enclose hazardous circuitry, nor is it part of safety or hazardous circuits; it is non-critical.
Enclosure/Casing | Assembly | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [1, 4] | 2 | 0.25 | The enclosure/casing is required by IEC 61010 and prevents access to hazardous circuitry; it is critical.
Tray | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | The tray is not indicated as enclosing, segregating, or preventing access to hazardous circuitry, nor is it required by IEC 61010 or used in safety/hazardous circuits. No evidence of special evaluation or abnormal operation reliance.
Rubber Feet | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | FALSE | 0.7 | [] | 0 | 0 | Rubber feet are not required by IEC 61010, do not enclose or segregate hazardous circuitry, and are not part of safety or hazardous circuits. No evidence of special evaluation or abnormal operation reliance.
Compressed Air Cylinder | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.95 | [8] | 1 | 0.125 | Compressed air cylinders present a potential hazard if changed or deleted, as pressurized systems are specifically addressed in IEC 61010 for mechanical hazards. Their presence and integrity are critical to prevent other hazards (e.g., explosion, rapid decompression).