## Sheet1
Component Name | Category | Confidence | Source Types | Evidence Count | Image URLs | Guide Reference | critical | confidence | rules_passed | rules_passed_count | rule_score | reasoning | photo_no
PRESSURE RELIEF SAFETY VALVE (125 PSI) | Safety Device | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 2, 3, 7, 8] | 5 | 0.625 | Pressure relief safety valves are required by IEC 61010 for overpressure protection in pneumatic/hydraulic systems. They are located in safety circuitry, directly prevent hazardous overpressure, and are relied upon for safe abnormal operation. Removal or failure would result in a significant hazard. | 1
FAN 1688-1554-ND | Cooling Fan | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [3, 7, 8] | 3 | 0.375 | Cooling fans are critical if they prevent overheating of hazardous circuitry, which may lead to fire or electric shock. Their removal or failure can result in unsafe abnormal operation and other hazards. | 1
MAC VALVE (VENT VALVE) | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [3, 4, 8] | 3 | 0.375 | A vent valve in a pneumatic system is located in hazardous circuitry and can prevent access to hazardous pressure. Its removal or failure could result in a hazardous condition, making it critical for safety compliance. | 1
NO VENT VALVE 126419 | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [3, 4, 8] | 3 | 0.375 | A vent valve in a pneumatic system can enclose or prevent access to hazardous pressure and is located in hazardous circuitry. Its removal or malfunction could result in uncontrolled release of pressure, presenting a safety hazard. | 1
E-REGULATOR ITV0050-3BL | Pneumatic Regulator | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [2, 3, 7] | 3 | 0.375 | E-regulator controls pressure in hazardous pneumatic circuits, is part of safety control, and is relied upon for safe abnormal operation (e.g., pressure regulation failure could cause hazard). | 1
NC VALVE 181168 | Valve | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [2, 3, 7] | 3 | 0.375 | NC (normally closed) valve is typically located in hazardous and safety-related pneumatic circuits, controlling access to pressurized gas and relied upon for safe abnormal operation (e.g., fail-safe shutoff). | 1
PRESSURE SENSOR SPT25-10-0200A | Sensor | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [2, 3, 7] | 3 | 0.375 | Pressure sensor is likely used in safety-related and hazardous pneumatic circuitry to monitor and control pressure, and may be relied upon for safe abnormal operation (e.g., overpressure detection). | 1
BACKPLANE | Mechanical/Electrical Interface | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.95 | [5, 8] | 2 | 0.25 | Backplane is relied upon for maintaining spacings and segregation of circuits, which is essential for electrical safety. Its removal or modification may introduce hazards. | 1
MECHANICAL PREREGULATOR | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.85 | [3, 7] | 2 | 0.25 | Mechanical preregulator is located in hazardous pneumatic circuitry and is relied upon for safe abnormal operation by limiting upstream pressure. | 1
MAIN MANIFOLD | Pneumatic Component | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.85 | [3, 8] | 2 | 0.25 | Main manifold is part of hazardous pneumatic circuitry; removal or modification could result in uncontrolled gas flow, presenting a hazard. | 1
PCB | Printed Circuit Board | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | PCB is likely to carry hazardous voltages and is relied upon to maintain required spacings and segregation per IEC 61010. Any failure or modification could compromise safety barriers. | 1
BASEPLATE AND SKINS | Mechanical Structure | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 0.9 | [4, 5] | 2 | 0.25 | Baseplate and skins typically form the enclosure, preventing access to hazardous parts and maintaining required spacings/segregation, satisfying rules 4 and 5. These are critical for compliance with IEC 61010. | 1
MAIN PCB | Printed Circuit Board | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | MAIN PCB is likely to contain hazardous circuitry and is relied upon for maintaining spacings/segregation per IEC 61010. Critical for electrical safety. | 1
POWER SUPPLY RPS-500-SF | Power Supply | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 7] | 3 | 0.375 | Power supply is required by standard, located in hazardous circuitry, and is relied upon for safe abnormal operation (e.g., overcurrent, overvoltage protection). | 1
POWER POWER ENTRY MODULE FPYSSA106FBR | Power Entry Module | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/CFE_block_diagram.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 4] | 3 | 0.375 | Power entry module is required by IEC 61010, located in hazardous mains circuitry, and encloses/prevents access to hazardous voltages. Its integrity is essential for user safety. | 1
Secondary Circuits (PCB) | Electronic Circuit | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | Secondary circuits on the PCB may be hazardous and are used to maintain spacings/segregation from primary circuits, as required by IEC 61010. | 2
DC 24V Output | Power Output | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 7] | 2 | 0.25 | The DC 24V output is part of hazardous circuitry if it can present hazardous energy, and is relied upon for safe abnormal operation (e.g., overvoltage, short-circuit protection). | 2
PRI (Primary Side) | Power Supply Section | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [3, 5] | 2 | 0.25 | The primary side of the power supply is hazardous (mains voltage) and is used to maintain spacings/segregation from secondary circuits, as required by IEC 61010. | 2
SWPS (RPS-500-24-C) | Power Supply | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 7] | 3 | 0.375 | The power supply is required by the standard, is located in hazardous circuitry (mains input), and is relied upon for safe abnormal operation (e.g., insulation, fault protection). | 2
AC Inlet | Connector | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 3, 4] | 3 | 0.375 | The AC inlet is required by IEC 61010 for mains connection, is located in hazardous (mains) circuitry, and encloses/prevents access to hazardous voltages. Its integrity is essential for user safety. | 2
Chassis Ground | Grounding | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 2, 7, 8] | 4 | 0.5 | Chassis ground is required by IEC 61010 for protective earthing, is part of safety circuitry, is relied upon for safe abnormal operation (faults), and its removal would create a shock hazard. | 2
Fuse | Protection Device | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 2, 3, 7] | 4 | 0.5 | The fuse is required by IEC 61010 for overcurrent protection, is part of safety circuitry, located in hazardous mains circuit, and is relied upon for safe abnormal operation (overcurrent/short circuit protection). | 2
Metal Enclosure | Enclosure | high | image | 1 | https://stintertekesusdev.blob.core.windows.net/cdr-test/PastedGraphic-1.png?sp=r&st=2025-12-17T13:00:21Z&se=2040-12-30T21:15:21Z&sv=2024-11-04&sr=c&sig=IvefAZb2x6KvlYtb22W5VK9DoQ2PDosWYtmtZO1pPCM%3D |  | TRUE | 1 | [1, 4] | 2 | 0.25 | The metal enclosure is required by IEC 61010 to prevent access to hazardous circuitry and provide protection against electric shock. | 2
Printed Circuit Board (PCB) | Electrical Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [3, 5] | 2 | 0.25 | PCB is located in hazardous circuitry and used to maintain spacings/segregation per IEC 61010 requirements. | guide
Ventilation Slots | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [4, 5] | 2 | 0.25 | Ventilation slots can prevent access to hazardous circuitry and maintain required spacings/segregation, satisfying rules 4 and 5 per IEC 61010. | guide
Rear Connections | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [3, 4] | 2 | 0.25 | Rear connections may provide access to hazardous voltages or circuits and often serve as entry points to hazardous areas, thus enclosure and prevention of access is required by IEC 61010. | guide
Enclosure/Casing | Assembly | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [1, 4] | 2 | 0.25 | The enclosure/casing is required by IEC 61010 and prevents access to hazardous circuitry; it is critical. | guide
Power Cord (C13 to 5-15) | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.95 | [1, 3, 4] | 3 | 0.375 | Power cord is required by IEC 61010, interfaces with hazardous voltage, and prevents direct access to hazardous circuitry. It is critical for electrical safety. | guide
Connection Lines and Cables | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [3, 5] | 2 | 0.25 | Connection lines and cables can carry hazardous voltages/currents and are relied upon to maintain required spacings and segregation between circuits per IEC 61010. | guide
Power Connection Port | Component | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.95 | [1, 3, 4] | 3 | 0.375 | Power connection port is required by IEC 61010, interfaces with hazardous voltage, and prevents direct access to hazardous circuitry. It is critical for electrical safety. | guide
Enclosure | Component | high | guide | 3 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 4] | 2 | 0.25 | Enclosure is required by IEC 61010 to prevent access to hazardous circuitry and is specifically referenced in the standard for protection against electric shock and mechanical hazards. | guide
Power Supply | Electrical Component | high | guide | 4 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 3, 7] | 3 | 0.375 | Power supply is required by IEC 61010, located in hazardous circuitry, and relied upon for safe abnormal operation (e.g., overvoltage, overcurrent protection). | guide
Power Switch | Component | high | guide | 2 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 1 | [1, 3, 4, 7] | 4 | 0.5 | Power switch is required by IEC 61010, is located in hazardous (mains) circuitry, encloses/prevents access to hazardous voltages, and is relied upon for safe abnormal operation (emergency shutoff). | guide
Sample Chamber | Fluidic Subsystem | high | guide | 2 |  | CellFE_Infinity_MTx_Operating_Manual.docx; CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.9 | [3, 4] | 2 | 0.25 | The Sample Chamber contains hazardous fluids and may enclose hazardous circuitry, satisfying rules 3 and 4. It is critical for user protection and containment. | guide
Infinity Instrument | Subsystem | medium | guide | 1 |  | CellFE_Infinity_MTx_Operating_Manual.docx | TRUE | 0.8 | [1, 3] | 2 | 0.25 | As a subsystem, the Infinity Instrument is required by the standard for evaluation and may contain hazardous circuitry; thus, it is critical. | guide