|  |  | IEC 61010-1 |  |  |
| --- | --- | --- | --- | --- |
| Clause | Requirement — Test |  | Result — Remark | Verdict |

| 6 | TABLE: Protection against electric shock - Block diagram of system Form A.5 |  |  |  |  |  |  |  |  |  | P |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pollution degree :2 |  |  |  |  | Overvoltage category: Ⅱ |  |  |  |  |  |  |
| Location or | Insulation type | Maximum working | CREEPAGE Distance (NOTE 3) |  |  |  |  | CLEARANCE (NOTE 3) | Test voltage | Comments |  |
| description | (NOTE 1) | voltage (NOTE 2) | PWB mm | CTI |  | Other mm | CTI | mm | (NOTE 2) V |  |  |
| A | BI | 240Vrms | 1.5 | ≧100 |  | 3.0 | >100 | 1.5 |  | Line to Neutral before fuse |  |
| B and C | BI | 240Vrms | 1.5 | ≧100 |  | 3.0 | >100 | 1.5 |  | Line neutral to GND |  |
| D | RI or DI | 240Vrms+24V dc | 3.0 | ≧100 |  | 6.0 | >100 | 3.0 |  | Part of power supply |  |
| E | BI | 24Vdc | 0.04 | ≧100 |  | 1.25 | ≧100 | 0.5 |  | Secondary to chassis |  |
| NOTE 1 – Type of insulation: NOTE 2 - Types of voltage NOTE 3 - INSTALLATION CATEGORIES BI = BASIC INSULATION Peak impulse test voltage (pulse) (OVERVOLTAGE CATEGORIES) DI = DOUBLE INSULATION r.m.s. or POLLUTION DEGREES which differ from PI = PROTECTIVE IMPEDANCE d.c. these should be shown under "Comments". RI = Reinforced INSULATION peak SI = Supplementary INSULATION |  |  |  |  |  |  |  |  |  |  |  |